Freescale Semiconductor, Inc.

Order this document

SEMICONDUCTOR TECHNICAL DATA

by MPX2202/D

| "# |  |
| --- | --- |
| ! " " |  |
| !!# ! ! |  |

The MPX2202/MPXV2202G device series is a silicon piezoresistive pressure sensor providing a highly accurate and linear voltage output — directly proportional to the applied pressure.  The sensor is a single monolithic silicon diaphragm with the strain gauge and a thin–film resistor network integrated on–chip. The chip is laser trimmed for precise span and offset calibration and temperature compensation.  They are designed for use in applications such as pump/motor controllers, robotics, level indicators, medical diagnostics, pressure switching, barometers, altimeters, etc.

Features

Motorola Preferred Device

0 to 200 kPa (0 to 29 psi)

40 mV FULL SCALE SPAN (TYPICAL)

   Temperature Compensated Over 0C to +85C

   Easy–to–Use Chip Carrier Package Options

   Available in Absolute, Differential and Gauge Con- figurations

Application Examples

   Pump/Motor Controllers

   Robotics

   Level Indicators

   Medical Diagnostics

   Pressure Switching

   Barometers

   Altimeters

Figure 1 illustrates a block diagram of the internal circuitry on the stand–alone pressure sensor chip.

+(

UNIBODY PACKAGE

MPX2202A/D CASE 344

SMALL OUTLINE PACKAGE SURFACE MOUNT

MPXV2202GP CASE 1369

( #( #

! " #)

) #  !"

) "% ' )*'

$"% #( ) $#

#

!  ' ) $#

' * )'.

+587

+587

MPX2202AP/GP CASE 344B

| MPXV2202DP CASE 1351 PIN NUMBER |  |  |  |
| --- | --- | --- | --- |
| 1 | Gnd | 5 | N/C |
| 2 | +Vout | 6 | N/C |
| 3 | VS | 7 | N/C |
| 4 | –Vout | 8 | N/C |

#

Figure 1. Temperature Compensated Pressure

Sensor Schematic

VOLTAGE OUTPUT versus

APPLIED DIFFERENTIAL PRESSURE

The differential voltage output of the sensor is

directly proportional to the differential pressure applied. The absolute sensor has a built–in reference vacu-

um. The output voltage will decrease as vacuum, relative to ambient, is drawn on the pressure (P1) side. The output voltage of the differential or gauge sensor

increases with increasing pressure applied to the

pressure (P1) side relative to the vacuum (P2) side. Similarly, output voltage increases as increasing vacu- um is applied to the vacuum (P2) side relative to the pressure (P1) side.

Preferred devices are Motorola recommended choices for future use and best overall value.

Replaces MPX2200/D

REV 2

MPX2202DP CASE 344C

MPX2202ASX/GSX CASE 344F

| NOTE: Pin 1 is noted by the notch in the lead. PIN NUMBER |  |  |  |
| --- | --- | --- | --- |
| 1 | Gnd | 3 | VS |
| 2 | +Vout | 4 | –Vout |

NOTE: Pin 1 is noted by the notch in the lead.

MPX2202GVP CASE 344D

Freescale

Semiconductor, Inc.

 Motorola, Inc. 2002

Motorola Sensor Device Data	1

For More Information On This Product,

F  r eescale Semiconductor, Inc.

MAXIMUM RATINGS(NOTE)

| Rating | Symbol | Value | Unit |
| --- | --- | --- | --- |
| Maximum Pressure (P1 > P2) | Pmax | 800 | kPa |
| Storage Temperature | Tstg | –40 to +125 | C |
| Operating Temperature | TA | –40 to +125 | C |

NOTE:   Exposure beyond the specified limits may cause permanent damage or degradation to the device.

OPERATING CHARACTERISTICS (VS = 10 Vdc, TA = 25C unless otherwise noted, P1 > P2)

| Characteristics | Symbol | Min | Typ | Max | Unit |
| --- | --- | --- | --- | --- | --- |
| Pressure Range(1) | POP | 0 | — | 200 | kPa |
| Supply Voltage | VS | — | 10 | 16 | Vdc |
| Supply Current | Io | — | 6.0 | — | mAdc |
| Full Scale Span(3) | VFSS | 38.5 | 40 | 41.5 | mV |
| Offset(4) | Voff | –1.0 | — | 1.0 | mV |
| Sensitivity | V/ P | — | 0.2 | — | mV/kPa |
| Linearity(5) MPX2202D Series MPX2202A Series | — | –0.6 –1.0 | — — | 0.4 1.0 | %VFSS |
| Pressure Hysteresis(5) (0 to 200 kPa) | — | — | 0.1 | — | %VFSS |
| Temperature Hysteresis(5) (–40C to +125C) | — | — | 0.5 | — | %VFSS |
| Temperature Effect on Full Scale Span(5) | TCVFSS | –2.0 | — | 2.0 | %VFSS |
| Temperature Effect on Offset(5) | TCVoff | –1.0 | — | 1.0 | mV |
| Input Impedance | Zin | 1000 | — | 2500 |  |
| Output Impedance | Zout | 1400 | — | 3000 |  |
| Response Time(6) (10% to 90%) | tR | — | 1.0 | — | ms |
| Warm–Up | — | — | 20 | — | ms |
| Offset Stability(7) | — | — | 0.5 | — | %VFSS |

NOTES:

1. 1.0 kPa (kiloPascal) equals 0.145 psi.

2. Device is ratiometric within this specified excitation range. Operating the device above the specified excitation range may induce additional error due to device self–heating.

3. Full Scale Span (VFSS) is defined as the algebraic difference between the output voltage at full rated pressure and the output voltage at the minimum rated pressure.

4. Offset (Voff) is defined as the output voltage at the minimum rated pressure.

5. Accuracy (error budget) consists of the following:

    Linearity:	Output deviation from a straight line relationship with pressure, using end point method, over the specified pressure range.

    Temperature Hysteresis:  Output deviation at any temperature within the operating temperature range, after the temperature is cycled to and from the minimum or maximum operating temperature points, with zero differential pressure applied.

    Pressure Hysteresis:	Output deviation at any pressure within the specified range, when this pressure is cycled to and from the minimum or maximum rated pressure, at 25C.

    TcSpan:	Output deviation at full rated pressure over the temperature range of 0 to 85C, relative to 25C.

    TcOffset:	Output deviation with minimum rated pressure applied, over the temperature range of 0 to 85C, relative to 25C.

6. Response Time is defined as the time for the incremental change in the output to go from 10% to 90% of its final value when subjected to a specified step change in pressure.

7. Offset stability is the product’s output deviation when subjected to 1000 hours of Pulsed Pressure, Temperature Cycling with Bias Test.

Freescale

Semiconductor, Inc.

2	For More Information On This Product,

Motorola Sensor Device Data

LINEARITY

Freescale Semiconductor,  In c  .

' ! ) +  +$!)   $*)%*)

Linearity refers to how well a transducer’s output follows the equation: Vout = Voff + sensitivity x P over the operating pressure range. There are two basic methods for calculating nonlinearity: (1) end point straight line fit (see Figure 2) or (2)

a least squares best line fit. While a least squares fit gives the “best case” linearity error (lower numerical value), the calculations required are burdensome.

Conversely, an end point fit will give the “worst case” error (often more desirable in error budget calculations) and the calculations are more straightforward for the user. Motorola’s specified pressure sensor linearities are based on the end point straight line method measured at the midrange pressure.

!  () (&* ' (  )

-    ' )

% ' $'" #

*'+

# %$ #)

()'     ) ! #  )

!  ()

(&* '

+ ) $#

()'    ) ! #

+ ) $#

$ ( )

%' ((*'      *!!(  !

Figure 2. Linearity Specification Comparison

ON–CHIP TEMPERATURE COMPENSATION and CALIBRATION

Figure 3 shows the output characteristics of the MPX2202/MPXV2202G series at 25C. The output is directly proportional to the differential pressure and is essentially a

straight line.

The effects of temperature on Full Scale Span and Offset are very small and are shown under Operating Characteristics.

$*)%*)  4+21

9

+(      +21

)      

%    %

" -

).%

" #

(% #

' #

).%

$  ( )

3%/

%(

%' ((*'

Figure 3. Output versus Pressure Differential

, '   $#

( ! $#    !

$ )

' #)  !   *

()  #! (( ()  !

" ) ! $+ '

%

%$-.

(

, '   $#

( ! $#    !

$ )

($!*)

%

()  #! (( ()  ! " ) ! $+ '

%$-.

(

!     ' "

' #)  !   *   ! " #)

%

$#

!     ' "

($!*)  ! " #)

%

$#

Figure 4. Cross–Sectional Diagrams (Not to Scale)

Figure 4 illustrates an absolute sensing die (right) and the differential or gauge die in the basic chip carrier (Case 344). A silicone gel isolates the die surface and wire bonds from the environment, while allowing the pressure signal to be transmitted to the silicon diaphragm.

The MPX2202/MPXV2202G series pressure sensor oper-

ating characteristics and internal reliability and qualification tests are based on use of dry air as the pressure media. Me- dia other than dry air may have adverse effects on sensor performance and long term reliability. Contact the factory for information regarding media compatibility in your application.

Freescale

Semiconductor,

Inc.

Motorola Sensor Device Data

For More Information On This Product,	3

F  r eescale Semiconductor, Inc.

PRESSURE (P1)/VACUUM (P2) SIDE IDENTIFICATION TABLE

Motorola designates the two sides of the pressure sensor as the Pressure (P1) side and the Vacuum (P2) side. The Pressure (P1) side is the side containing the silicone gel which isolates the die from the environment. The differential or gauge sensor is designed to operate with positive differen-

tial pressure applied, P1 > P2. The absolute sensor is designed for vacuum applied to P1 side.

The Pressure (P1) side may be identified by using the

table below:

| Part Number |  | Case Type | Pressure (P1) Side Identifier |
| --- | --- | --- | --- |
| MPX2202A | MPX2202D | 344 | Stainless Steel Cap |
| MPX2202DP |  | 344C | Side with Part Marking |
| MPX2202AP | MPX2202GP | 344B | Side with Port Attached |
| MPX2202GVP |  | 344D | Stainless Steel Cap |
| MPX2202ASX | MPX2202GSX | 344F | Side with Port Attached |
| MPXV2202GP |  | 1369 | Side with Port Attached |
| MPXV2202DP |  | 1351 | Side with Part Marking |

ORDERING INFORMATION — UNIBODY PACKAGE (MPX2202 SERIES)

| Device Type | Options | Case Type | MPX Series |  |
| --- | --- | --- | --- | --- |
|  |  |  | Order Number | Device Marking |
| Basic Element | Absolute, Differential | 344 | MPX2202A MPX2202D | MPX2202A MPX2202D |
| Ported Elements | Differential, Dual Port | 344C | MPX2202DP | MPX2202DP |
|  | Absolute, Gauge | 344B | MPX2202AP MPX2202GP | MPX2202AP MPX2202GP |
|  | Absolute, Gauge Axial | 344F | MPX2202ASX MPX2202GSX | MPX2202A MPX2202D |
|  | Gauge, Vacuum | 344D | MPX2202GVP | MPX2202GVP |

ORDERING INFORMATION — SMALL OUTLINE PACKAGE (MPXV2202G SERIES)

| Device Type | Options | Case No. | MPX Series Order No. | Packing Options | Marking |
| --- | --- | --- | --- | --- | --- |
| Ported Elements | Gauge, Side Port, SMT | 1369 | MPXV2202GP | Trays | MPXV2202G |
|  | Differential, Dual Port, SMT | 1351 | MPXV2202DP | Trays | MPXV2202G |

Freescale

Semiconductor, Inc.

4	For More Information On This Product,

Motorola Sensor Device Data

Freescale Semiconductor,  In c  .

SMALL OUTLINE PACKAGE DIMENSIONS

A	E

D

e

e/2

.014 (0.35)

L	A1

DETAIL G

F

B	E1

| b |  | " |  |  |  |
| --- | --- | --- | --- | --- | --- |

#$) (

$#)'$!! #    " #( $#   #

#) '%' )  " #( $#( # )$! ' #  ( % '

("  .    "

" #( $#(      #        $ #$)  # !*   "$!

! ( $' %'$)'*( $#( "$!   ! ( $'

%'$)'*( $#( (  !! #$)  -

% ' (

" #( $# 0   $ ( #$) # !*     " '

| %'$)'*( $# !!$, ! " ' %'$)'*( $# ( !! " - "*" |  |  |  |  |
| --- | --- | --- | --- | --- |
| ! |  |  |  |  |
|  | ( |  | ( |  |
| " | ( |  | ( |  |
|  |  |  |  |  |

 T	N

K

A	M

P

DETAIL G

C

CASE 1369–01

ISSUE O

Freescale

Semiconductor, Inc.

Motorola Sensor Device Data

For More Information On This Product,	5

F  r eescale Semiconductor, Inc.

SMALL OUTLINE PACKAGE DIMENSIONS—CONTINUED

A	E

e

e/2

D

.014 (0.35)

L	A1

DETAIL G

F

B	E1

| b |  | " |  |  |  |
| --- | --- | --- | --- | --- | --- |

().!

% #      #

+587

+6

+587

#

#

#

#

().!

% #     #

+6

#

+587

#

#

#

#

#$) (

$#)'$!! #    " #( $#   #

#) '%' )  " #( $#( # )$! ' #  ( % '

("  .    "

" #( $#(      #        $ #$)  # !*   "$!

! ( $' %'$)'*( $#( "$!   ! ( $'

%'$)'*( $#( (  !! #$)  -

% ' (

" #( $# 0   $ ( #$) # !*     " '

| %'$)'*( $# !!$, ! " ' %'$)'*( $# ( !! " - "*" |  |  |  |  |
| --- | --- | --- | --- | --- |
| ! |  |  |  |  |
| " | ( |  | ( |  |
|  |  |  |  |  |

N	 T

A	M

P		K DETAIL G

C

CASE 1351–01

ISSUE O

Freescale

Semiconductor, Inc.

6	For More Information On This Product,

Motorola Sensor Device Data

Freescale Semiconductor,  In c  .

UNIBODY PACKAGE DIMENSIONS

#$) (

C	      " #( $# #   #  )$! ' #  #  % '  ("

.    "

R	     $#)'$!! #    " #( $#   #

" #( $#      ( # !*( + $ )   "$! ()$% ' #   "$!  ()$% ' #  #$) )$  -

M	Z

B	–A–

N

PIN 1	L

–T–

F J	G

F	Y

| D 4 PL |  | " | ) | " |
| --- | --- | --- | --- | --- |

DAMBAR TRIM ZONE: THIS IS INCLUDED WITHIN DIM. “F” 8 PL

().!

% #      '$*#

$*)%*)

(*%%!.

$*)%*)

().!

% #     +

(*%%!.

(*%%!.

'$*#

().!

| % # # |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | ( |  | ( |  |
|  | #$" |  | #$" |  |

+$*)

+(

+$*)

CASE 344–15

ISSUE Z

–T–

R

–A– U

L

H

#$) (

" #( $# #   #  )$! ' #  #  % '  #(

.

| $#)'$!! # " #( $# # |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | ( |  | ( |  |
|  | ( |  | ( |  |

N	PORT #1

POSITIVE PRESSURE (P1)

B

–Q–

–P–

PIN 1	K

"

J

C

) &   (	S

F

G

D 4 PL

"

) (  (   &  (

().!

% #      '$*#

$*)%*)

(*%%!.

$*)%*)

CASE 344B–01

ISSUE B

Freescale

Semiconductor, Inc.

Motorola Sensor Device Data

For More Information On This Product,	7

F  r eescale Semiconductor, Inc.

UNIBODY PACKAGE DIMENSIONS — CONTINUED

V R

PORT #2

PORT #1

–A– U

W	L

H

#$) (

" #( $# #   #  )$! ' #  #  % '  #(

.    "

| $#)'$!! # " #( $# # |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | ( |  | ( |  |
|  | ( |  | ( |  |

PORT #2

N	VACUUM

PORT #1

POSITIVE PRESSURE

(P2)	(P1)

B

–Q–

–T–	–T– J

C

–P–

"

K

) &   (	S F

G

D 4 PL

|  |  | " | ) | ( ( | & ( |
| --- | --- | --- | --- | --- | --- |

CASE 344C–01

ISSUE B

().!

% #      '$*#

$*)%*)

(*%%!.

$*)%*)

–T– R

PORT #2

–A– U

L

H

#$) (

" #( $# #   #  )$! ' #  #  % '  #(

.

| $#)'$!! # " #( $# # |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | ( |  | ( |  |
|  | ( |  | ( |  |

VACUUM	POSITIVE

(P2)

N

PRESSURE (P1)

–Q–

B

C

–P– J

"

PIN 1

F

) &   (

K S

G

| D 4 PL |  | " | ) | ( ( | & ( |
| --- | --- | --- | --- | --- | --- |

().!

% #      '$*#

$*)%*)

(*%%!.

$*)%*)

CASE 344D–01

ISSUE B

Freescale

Semiconductor, Inc.

8	For More Information On This Product,

Motorola Sensor Device Data

Freescale Semiconductor,  In c  .

UNIBODY PACKAGE DIMENSIONS — CONTINUED

–T– C

E

A

–Q– U

#$) (

" #( $# #  #  )$! ' #  # % '

#(  .    "

| $#)'$!! # " #( $# # |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | ( |  | ( |  |

V

PORT #1

POSITIVE PRESSURE (P1)

J

–P–

"

) &   "

F

G

N		B R

PIN 1

S K

D 4 PL

"

) %    (   &  (

().!

% #      '$*#

+     $*)

+ (*%%!.

+     $*)

CASE 344F–01

ISSUE B

Freescale

Semiconductor, Inc.

Motorola Sensor Device Data

For More Information On This Product,	9

F  r eescale Semiconductor, Inc.

NOTES

Freescale

Semiconductor, Inc.

10	For More Information On This Product,

Motorola Sensor Device Data

Freescale Semiconductor,  In c  .

NOTES

Freescale

Semiconductor, Inc.

Motorola Sensor Device Data

For More Information On This Product,	11

F  r eescale Semiconductor, Inc.

Motorola reserves the right to make changes without further notice to any products herein. Motorola makes no warranty, representation or guarantee regarding the suitability of its products for any particular purpose, nor does Motorola assume any liability arising out of the application or use of any product or circuit, and specifically disclaims any and all liability, including without limitation consequential or incidental damages. “Typical” parameters can and do vary in different applications and actual performance may vary over time. All operating parameters, including “Typicals” must be validated for each customer application by customer’s technical experts.  Motorola does not convey any license under its patent rights nor the rights of others.  Motorola products are not designed, intended, or authorized for use as components in systems intended for surgical implant into the body, or other applications intended to support or sustain life, or for any other application in which the failure of the Motorola product could create a situation where personal injury or death may occur. Should Buyer purchase or use Motorola products for any such unintended or unauthorized application, Buyer shall indemnify and hold Motorola and its  officers, employees, subsidiaries, affiliates, and distributors harmless against all claims, costs, damages, and expenses, and reasonable attorney fees arising out of, directly or indirectly, any claim of personal injury or death associated with such unintended or unauthorized use, even if such claim alleges that Motorola was negligent regarding the design or manufacture of the part.  Motorola and the Stylized M Logo are registered trademarks of Motorola, Inc.  Motorola, Inc. is an Equal Opportunity/Affirmative Action Employer.

MOTOROLA and the Stylized M Logo are registered in the US Patent & Trademark Office. All other product or service names are the property of their respective owners.

Motorola, Inc. 2002.

How to reach us:

USA/EUROPE/Locations Not Listed: Motorola Literature Distribution; P.O. Box 5405, Denver, Colorado 80217. 1–303–675–2140 or 1–800–441–2447

JAPAN: Motorola Japan Ltd.; SPS, Technical Information Center, 3–20–1, Minami–Azabu. Minato–ku, Tokyo 106–8573 Japan. 81–3–3440–3569

ASIA/PACIFIC: Motorola Semiconductors H.K. Ltd.; Silicon Harbour Centre, 2 Dai King Street, Tai Po Industrial Estate, Tai Po, N.T., Hong Kong. 852–26668334

Technical Information Center: 1–800–521–6274

HOME PAGE: http://www.motorola.com/semiconductors/

Freescale

Semiconductor, Inc.

For More Information On This Product,

12	Motorola Sensor Device Data

MPX2202/D