Freescale Semiconductor

50 kPa Uncompensated

Silicon Pressure Sensors

The MPX53 series silicon piezoresistive pressure sensors provide a very accurate and linear voltage output, directly proportional to the applied pressure. These standard, low cost, uncompensated sensors permit manufacturers to design and add their own external temperature compensating and signal conditioning networks. Compensation techniques are simplified because of the predictability of Freescale's single element strain gauge design.

Features

• Low Cost

• Patented Silicon Shear Stress Strain Gauge Design

• Ratiometric to Supply Voltage

• Easy to Use Chip Carrier Package Options

• 60 mV Span (Typical)

• Differential and Gauge Options

MPX53

Rev 7, 05/2009

MPX53

Series

0 to 50 kPa (0 to 7.25 psi)

60 mV Full Scale Span

(Typical)

Application Examples

•   Air Movement Control

•   Environmental Control Systems

•   Level Indicators

•   Leak Detection

•   Medical Instrumentation

•   Industrial Controls

•   Pneumatic Control Systems

•   Robotics

| Marking ORDERING INFORMATION |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Device Name | Package Options | Case No. | # of Ports |  |  | Pressure Type |  |  | Device |
|  |  |  | None | Single | Dual | Gauge | Differential | Absolute |  |
| Unibody Package (MPX53 Series) |  |  |  |  |  |  |  |  |  |
| MPX53D | Tape & Reel | 344 | • |  |  |  | • |  | MPX53D |
| MPX53DP | Rail | 344C |  |  | • |  | • |  | MPX53DP |
| MPX53GP | Rail | 344B |  | • |  | • |  |  | MPX53GP |
| Small Outline Package (MPXV53G Series) |  |  |  |  |  |  |  |  |  |
| MPXV53GC7U | Rail | 482C |  | • |  | • |  |  | MPXV53G |

SMALL OUTLINE PACKAGE

UNIBODY PACKAGES

MPXV53GC7U CASE 482C-03

MPX53D CASE 344-15

MPX53GP CASE 344B-01

MPX53DP CASE 344C-01

© Freescale Semiconductor, Inc., 2007-2009. All rights reserved.

Operating Characteristics

Table 1. Operating Characteristics (VS = 3.0 Vdc, TA = 25°C unless otherwise noted, P1 > P2)

| Characteristic | Symbol | Min | Typ | Max | Units |
| --- | --- | --- | --- | --- | --- |
| Pressure Range(1) | POP | 0 | — | 50 | kPa |
| Supply Voltage(2) | VS | — | 3.0 | 6.0 | VDC |
| Supply Current | IO | — | 6.0 | — | mAdc |
| Full Scale Span(3) | VFSS | 45 | 60 | 90 | mV |
| Offset(4) | VOFF | 0 | 20 | 35 | mV |
| Sensitivity | ΔV/ΔΡ | — | 1.2 | — | mV/kPa |
| Linearity | — | –0.6 | — | 0.4 | %VFSS |
| Pressure Hysteresis (0 to 50 kPa) | — | — | ±0.1 | — | %VFSS |
| Temperature Hysteresis | — | — | ±0.5 | — | %VFSS |
| Temperature Coefficient of Full Scale Span | TCVFSS | –0.22 | — | -0.16 | %VFSS/°C |
| Temperature Coefficient of Offset | TCVOFF | — | ±15 | — | µV/°C |
| Temperature Coefficient of Resistance | TCR | 0.21 | — | 0.27 | %ZIN/°C |
| Input Impedance | ZIN | 355 | — | 505 | Ω |
| Output Impedance | ZOUT | 750 | — | 1875 | Ω |
| Response Time(5) (10% to 90%) | tR | — | 1.0 | — | ms |
| Warm-Up Time(6) | — | — | 20 | — | ms |
| Offset Stability(7) | — | — | ±0.5 | — | %VFSS |

1.  1.0 kPa (kiloPascal) equals 0.145 psi.

2.  Device is ratiometric within this specified excitation range. Operating the device above the specified excitation range may induce additional error due to device self-heating.

3.  Full Scale Span (VFSS) is defined as the algebraic difference between the output voltage at full rated pressure and the output voltage at the minimum rated pressure.

4.  Offset (VOFF) is defined as the output voltage at the minimum rated pressure.

5.  Response Time is defined as the time for the incremental change in the output to go from 10% to 90% of its final value when subjected to a specified step change in pressure.

6.  Warm-up Time is defined as the time required for the product to meet the specified output voltage after the pressure is stabilized.

7.  Offset stability is the product’s output deviation when subjected to 1000 hours of Pulsed Pressure, Temperature Cycling with Bias Test.

MPX53

Maximum Ratings

Table 2. Maximum Ratings(1)

| Rating | Symbol | Value | Unit |
| --- | --- | --- | --- |
| Maximum Pressure (P1 > P2) | PMAX | 175 | kPa |
| Burst Pressure (P1 > P2) | PBurst | 200 | kPa |
| Storage Temperature | TSTG | –40 to +125 | °C |
| Operating Temperature | TA | – 40 to +125 | °C |

1.   Exposure beyond the specified limits may cause permanent damage or degradation to the device.

Figure 1 shows a schematic of the internal circuitry on the stand-alone pressure sensor chip.

3

+VS

Sensor

2

+VOUT

4

-VOUT

1

GND

Figure 1. Uncompensated Pressure Sensor Schematic

Voltage Output versus Applied Differential Pressure

The differential voltage output of the sensor is directly proportional to the differential pressure (P1) relative to the vacuum side (P2). Similarly, output voltage increases as

increasing vacuum is applied to the vacuum side (P2) relative to the pressure side (P1).

MPX53

Temperature Compensation

Figure 2 shows the typical output characteristics of the

MPX53 series over temperature.

The piezoresistive pressure sensor element is a semiconductor device which gives an electrical output signal proportional to the pressure applied to the device. This device uses a unique transverse voltage diffused semiconductor strain gauge which is sensitive to stresses produced in a thin silicon diaphragm by the applied pressure.

Because this strain gauge is an integral part of the silicon diaphragm, there are no temperature effects due to differences in the thermal expansion of the strain gauge and the diaphragm, as are often encountered in bonded strain gauge pressure sensors. However, the properties of the strain gauge itself are temperature dependent, requiring that the device be temperature compensated if it is to be used over an extensive temperature range.

Temperature compensation and offset calibration can be achieved rather simply with additional resistive components, or by designing your system using the MPX2053 series sensors.

Several approaches to external temperature

compensation over –40 to +125°C and 0 to + 80°C are presented in Freescale Application Note, AN840.

100

LINEARITY

Linearity refers to how well a transducer's output follows the equation: Vout = Voff + (sensitivity x P) over the operating pressure range (see Figure 3). There are two basic methods

for calculating nonlinearity: (1) end point straight line fit or

(2) a least squares best line fit. While a least squares fit gives the “best case” linearity error (lower numerical value), the calculations required are burdensome.

Conversely, an end point fit will give the “worst case” error (often more desirable in error budget calculations) and the calculations are more straightforward for the user. Freescale’s specified pressure sensor linearities are based on the end point straight line method measured at the midrange pressure.

Figure 4 illustrates the differential or gauge configuration in the unibody chip carrier (Case 344). A silicone gel isolates the die surface and wire bonds from the environment, while allowing the pressure signal to be transmitted to the silicon diaphragm.

The MPX53 series pressure sensor operating characteristics and internal reliability and qualification tests are based on use of dry air as the pressure media. Media other than dry air may have adverse effects on sensor performance and long term reliability. Refer to application note AN3728, for more information regarding media compatibility.

90	MPX53

80	VS = 3 Vdc

70	P1 > P2

Output (mVdc)

60

50

40

30

20

10

0

–40°C

+25°C

+ 125ºC

Span Range (Typ)

Offset

(Typ)

PSI

kPa

0	1	2	3	4	5	6	7	8

0	10	20	30	40	50

Pressure Differential

Figure 2. Output vs. Pressure Differential

90

80

70

60	Actual

50

Output (mVdc)

40

Linearity

Span

(VFSS)

Wire Bond

Silicone

Die Coat 	Die

P1

Stainless Steel

Metal Cover

Epoxy

Case

30	Theoretical

Lead Frame 	RTV Die

Bond

20

10

0

0	Pressure (kPa)

P2

Offset

(VOFF)

MAX	POP

Figure 3. Linearity Specification Comparison	Figure 4. Unibody Package — Cross-Sectional Diagram

(Not to Scale)

MPX53

PRESSURE (P1)/VACUUM (P2) SIDE IDENTIFICATION TABLE

| Freescale designates the two sides of the pressure sensor as the Pressure (P1) side and the Vacuum (P2) side. The Pressure (P1) side is the side containing silicone gel which isolates the die from the environment. The Freescale MPX pressure sensor is designed to operate with positive differential pressure applied, P1 > P2. Part Number | Case Type | Pressure (P1) Side Identifier |
| --- | --- | --- |
| MPX53D | 344 | Stainless Steep Cap |
| MPX53DP | 344C | Side with Port Marking |
| MPX53GP | 344B | Side with Port Attached |
| MPXV53 Series | 482C | Side with Port Attached |

The Pressure (P1) side may be identified by using the following table.

MPX53

PACKAGE DIMENSIONS

C

R

M

1	2	3	4

NOTES:

1. DIMENSIONINGANDTOLERANCINGPERASME Y14.5M, 1994.

2. CONTROLLINGDIMENSION: INCH.

3. DIMENSION-A- ISINCLUSIVEOFTHEMOLD STOPRING. MOLDSTOPRINGNOTTOEXCEED

| Z 16.00 (0.630). DIM | INCHES |  | MILLIMETERS |  |
| --- | --- | --- | --- | --- |
|  | MIN | MAX | MIN | MAX |
| A | 0.595 | 0.630 | 15.11 | 16.00 |
| B | 0.514 | 0.534 | 13.06 | 13.56 |
| C | 0.200 | 0.220 | 5.08 | 5.59 |
| D | 0.016 | 0.020 | 0.41 | 0.51 |
| F | 0.048 | 0.064 | 1.22 | 1.63 |
| G | 0.100 BSC |  | 2.54 BSC |  |
| J | 0.014 | 0.016 | 0.36 | 0.40 |
| L | 0.695 | 0.725 | 17.65 | 18.42 |
| M | 30˚ NOM |  | 30˚ NOM |  |
| N | 0.475 | 0.495 | 12.07 | 12.57 |
| R | 0.430 | 0.450 | 10.92 | 11.43 |
| Y | 0.048 | 0.052 | 1.22 | 1.32 |
| Z | 0.106 | 0.118 | 2.68 | 3.00 |

B	-A-

PIN 1

-T-

SEATING PLANE

J

F

N

1   2  3   4	L

G	F

Y

| D 4 PL |  | 0.136 (0.005) M | T | A M |
| --- | --- | --- | --- | --- |

DAMBAR TRIM ZONE: THIS IS INCLUDED WITHIN DIM. "F" 8 PL

CASE 344-15

ISSUE AA UNIBODY PACKAGE

SEATING PLANE

-T-

R

-A- U

L

H

NOTES:

1. DIMENSIONINGANDTOLERANCINGPERANSI Y14.5M, 1982.

| 2. CONTROLLINGDIMENSION: INCH. DIM | INCHES |  | MILLIMETERS |  |
| --- | --- | --- | --- | --- |
|  | MIN | MAX | MIN | MAX |
| A | 1.145 | 1.175 | 29.08 | 29.85 |
| B | 0.685 | 0.715 | 17.40 | 18.16 |
| C | 0.305 | 0.325 | 7.75 | 8.26 |
| D | 0.016 | 0.020 | 0.41 | 0.51 |
| F | 0.048 | 0.064 | 1.22 | 1.63 |
| G | 0.100 BSC |  | 2.54 BSC |  |
| H | 0.182 | 0.194 | 4.62 | 4.93 |
| J | 0.014 | 0.016 | 0.36 | 0.41 |
| K | 0.695 | 0.725 | 17.65 | 18.42 |
| L | 0.290 | 0.300 | 7.37 | 7.62 |
| N | 0.420 | 0.440 | 10.67 | 11.18 |
| P | 0.153 | 0.159 | 3.89 | 4.04 |
| Q | 0.153 | 0.159 | 3.89 | 4.04 |
| R | 0.230 | 0.250 | 5.84 | 6.35 |
| S | 0.220 | 0.240 | 5.59 | 6.10 |
| U | 0.910 BSC |  | 23.11 BSC |  |

N	PORT #1

POSITIVE PRESSURE (P1)

B

-Q-

-P- J

C

0.25 (0.010) M

PIN 1 	K

T  Q  S		S F

G

D 4 PL

0.13 (0.005) M

T  S S     Q S

STYLE1:

PIN1. GROUND

2. + OUTPUT

3. + SUPPLY

4  - OUTPUT

CASE 344B-01

ISSUE B UNIBODY PACKAGE

MPX53

PACKAGE DIMENSIONS

V R

PORT #2

PORT #1

-A- U

W	L

H

NOTES:

1. DIMENSIONINGANDTOLERANCINGPERANSI Y14.5M, 1982.

| 2. CONTROLLINGDIMENSION: INCH. DIM | INCHES |  | MILLIMETERS |  |
| --- | --- | --- | --- | --- |
|  | MIN | MAX | MIN | MAX |
| A | 1.145 | 1.175 | 29.08 | 29.85 |
| B | 0.685 | 0.715 | 17.40 | 18.16 |
| C | 0.405 | 0.435 | 10.29 | 11.05 |
| D | 0.016 | 0.020 | 0.41 | 0.51 |
| F | 0.048 | 0.064 | 1.22 | 1.63 |
| G | 0.100 BSC |  | 2.54 BSC |  |
| H | 0.182 | 0.194 | 4.62 | 4.93 |
| J | 0.014 | 0.016 | 0.36 | 0.41 |
| K | 0.695 | 0.725 | 17.65 | 18.42 |
| L | 0.290 | 0.300 | 7.37 | 7.62 |
| N | 0.420 | 0.440 | 10.67 | 11.18 |
| P | 0.153 | 0.159 | 3.89 | 4.04 |
| Q | 0.153 | 0.159 | 3.89 | 4.04 |
| R | 0.063 | 0.083 | 1.60 | 2.11 |
| S | 0.220 | 0.240 | 5.59 | 6.10 |
| U | 0.910 BSC |  | 23.11 BSC |  |
| V | 0.248 | 0.278 | 6.30 | 7.06 |
| W | 0.310 | 0.330 | 7.87 | 8.38 |

PORT #2

VACUUM

PORT #1

POSITIVE PRESSURE

N	(P2)	(P1)

-Q-

SEATING PLANE

B	SEATING PLANE

-P-	K

-T-	-T-

J

C

0.25 (0.010) M

T  Q S	S F

G

D 4 PL

|  |  | 0.13 (0.005) M | T | S S | Q S |
| --- | --- | --- | --- | --- | --- |

CASE 344C-01

STYLE1:

PIN1. GROUND

ISSUE B UNIBODY PACKAGE

CASE 482C-03

ISSUE B

SMALL OUTLINE PACKAGE

MPX53

How to Reach Us:

Home Page:

www.freescale.com

Web Support:

http://www.freescale.com/support

USA/Europe or Locations Not Listed:

Freescale Semiconductor, Inc. Technical Information Center, EL516

2100 East Elliot Road

Tempe, Arizona 85284

1-800-521-6274 or +1-480-768-2130 www.freescale.com/support

Europe, Middle East, and Africa: Freescale Halbleiter Deutschland GmbH Technical Information Center Schatzbogen 7

81829 Muenchen, Germany

+44 1296 380 456 (English)

+46 8 52200080 (English)

+49 89 92103 559 (German)

+33 1 69 35 48 48 (French)

www.freescale.com/support

Japan:

Freescale Semiconductor Japan Ltd. Headquarters

ARCO Tower 15F

1-8-1, Shimo-Meguro, Meguro-ku, Tokyo 153-0064

Japan

0120 191014 or +81 3 5437 9125 support.japan@freescale.com

Asia/Pacific:

Freescale Semiconductor China Ltd. Exchange Building 23F

No. 118 Jianguo Road Chaoyang District Beijing 100022

China

+86 010 5879 8000 support.asia@freescale.com

For Literature Requests Only:

Freescale Semiconductor Literature Distribution Center

1-800-441-2447 or +1-303-675-2140

Fax: +1-303-675-2150

LDCForFreescaleSemiconductor@hibbertgroup.com

Information in this document is provided solely to enable system and software implementers to use Freescale Semiconductor products. There are no express or implied copyright licenses granted hereunder to design or fabricate any integrated circuits or integrated circuits based on the information in this document.

Freescale Semiconductor reserves the right to make changes without further notice to any products herein. Freescale Semiconductor makes no warranty, representation or guarantee regarding the suitability of its products for any particular purpose, nor does Freescale Semiconductor assume any liability arising out of the application or use of any product or circuit, and specifically disclaims any and all liability, including without limitation consequential or incidental damages. “Typical” parameters that may be provided in Freescale Semiconductor data sheets and/or specifications can and do vary in different applications and actual performance may vary over time. All operating parameters, including “Typicals”, must be validated for each customer application by customer’s technical experts. Freescale Semiconductor does not convey any license under its patent rights nor the rights of others. Freescale Semiconductor products are not designed, intended, or authorized for use as components in systems intended for surgical implant into the body, or other applications intended to support or sustain life, or for any other application in which the failure of the Freescale Semiconductor product could create a situation where personal injury or death may occur. Should Buyer purchase or use Freescale Semiconductor products for any such unintended or unauthorized application, Buyer shall indemnify and hold Freescale Semiconductor and its officers, employees, subsidiaries, affiliates, and distributors harmless against all claims, costs, damages, and expenses, and reasonable attorney fees arising out of, directly or indirectly, any claim of personal injury or death associated with such unintended or unauthorized use, even if such claim alleges that Freescale Semiconductor was negligent regarding the design or manufacture of the part.

Freescale™ and the Freescale logo are trademarks of Freescale Semiconductor, Inc. All other product or service names are the property of their respective owners.

© Freescale Semiconductor, Inc. 2009. All rights reserved.

MPX53

Rev. 7

05/2009