19-4646; Rev 2; 6/09

High-Temperature Logger iButton with 8KB Data-Log Memory

General Description The DS1922E temperature logger iButton® is a rugged, self-sufficient system that measures temperature and records the result in a protected  memory section. The recording is done at a user-defined rate. A total of 8192

8-bit readings or 4096 16-bit readings, taken at equidis-

tant intervals ranging from 1s to 273hr, can be stored. Additionally, 576 bytes of SRAM store application-spe- cific information. A mission to collect data can be pro- grammed  to begin  immediately,  after a user-defined delay, or after a temperature  alarm. Access  to the memory and control functions can be password  pro- tected. The DS1922E is configured  and communicates with a host-computing device through the serial 1-Wire® protocol, which requires only a single data lead and a ground return. Every DS1922E is factory lasered with a guaranteed  unique 64-bit registration  number that allows for absolute traceability.  The durable  stainless- steel package  is highly resistant to environmental haz- ards such as dirt, moisture, and shock.

Applications High-Temperature Logging (Process Monitoring, Industrial Temperature Monitoring)

Steam Sterilization

Features

♦ Automatically Wakes Up, Measures Temperature, and Stores Values in 8KB of Data-Log Memory in

8- or 16-Bit Format

♦ Digital Thermometer Measures Temperature with

8-Bit (0.5°C) or 11-Bit (0.0625°C) Resolution

♦ Temperature Accuracy: ±1.5°C from +110°C to

+140°C, ±7°C typical from +15°C to +110°C

♦ Water Resistant or Waterproof if Placed Inside DS9107 iButton Capsule (Exceeds Water Resistant 3 ATM Requirements)

♦ Sampling Rate from 1s Up to 273hr

♦ Programmable High and Low Trip Points for

Temperature Alarms

♦ Programmable Recording Start Delay After Elapsed

Time or Upon a Temperature Alarm Trip Point

♦ Quick Access to Alarmed Devices Through 1-Wire

Conditional Search Function

♦ 576 Bytes of General-Purpose Memory

♦ Two-Level Password Protection of All Memory and Configuration Registers

♦ Communicates to Host with a Single Digital Signal

Up to 15.4kbps at Standard Speed or Up to

125kbps in Overdrive Mode Using 1-Wire Protocol

♦ Operating Temperature Range: +15°C to +140°C

Common iButton Features

♦ Digital Identification and Information by

Momentary Contact

♦ Unique Factory-Lasered 64-Bit Registration Number Ensures Error-Free Device Selection and Absolute Traceability Because No Two Parts Are Alike

♦ Built-In Multidrop Controller for 1-Wire Net

♦ Chip-Based Data Carrier Compactly Stores

Information

♦ Data Can Be Accessed While Affixed to Object

♦ Button Shape is Self-Aligning with Cup-Shaped

Probes

♦ Durable Stainless-Steel Case Engraved with Registration Number Withstands Harsh Environments

♦ Easily Affixed with Self-Stick Adhesive Backing, Latched by Its Flange, or Locked with a Ring Pressed Onto Its Rim

♦ Presence Detector Acknowledges When Reader

First Applies Voltage

♦ Meets UL 913 (4th Edit.); Intrinsically Safe Apparatus: Approved Under Entity Concept for Use in Class I, Division 1, Group A, B, C, and D Locations*

Ordering Information

PART 	TEMP RANGE 	PIN-PACKAGE

DS1922E-F5# 	+15°C to +140°C 	F5 iButton

# Denotes a RoHS-compliant device that may include lead(Pb)

that is exempt under the RoHS requirements.

Examples of Accessories

| PART | ACCESSORY |
| --- | --- |
| DS9093RA | Mount ing Lock Ring |
| DS9107 | iButton Capsule |
| DS9490B | USB to 1-Wire Adapter |

Pin Configuration appears at end of data sheet.

DS1922E

iButton and 1-Wire are registered trademarks of Maxim

Integrated Products, Inc.

*Application pending.

ABSOLUTE MAXIMUM RATINGS

I/O Voltage Range to GND .......................................-0.3V to +6V I/O Sink Current...................................................................20mA Operating  Temperature  Range ........................+15°C to +140°C

Junction Temperature ......................................................+150°C Storage Temperature  Range............................-25°C to +140°C*

DS1922E

*Storage or operation above +50°C significantly reduces battery life with an upper limit of 300hr cumulative at +140°C. The recom- mended storage temperature for maximum battery lifetime is between +5°C and +35°C.

Stresses beyond those listed under “Absolute Maximum Ratings” may cause permanent damage to the device. These are stress ratings only, and functional operation of the device at these or any other conditions  beyond those indicated  in the operational sections of the specifications  is not implied. Exposure to absolute maximum rating conditions for extended periods may affect device reliability.

ELECTRICAL CHARACTERISTICS

(VPUP = 3.0V to 5.25V.)

| PARAMETER | SYMBOL | CONDITIONS | MIN TYP MAX | UNITS |
| --- | --- | --- | --- | --- |
| Operating Temperature | TA | DS1922E (Note 1) | +15 +140 | °C |
| I/O PIN GENERAL DATA |  |  |  |  |
| 1-Wire Pullup Resistance | RPUP | (Notes 2, 3) | 2.2 | k |
| Input Capacitance | CIO | (Note 4) | 100 800 | pF |
| Input Load Current | IL | I/O pin at VPUP | 6 10 | μA |
| High-to-Low Switching Threshold | VTL | (Notes 5, 6) | 0.4 3.2 | V |
| Input Low Voltage | VIL | (Notes 2, 7) | 0.3 | V |
| Low-to-High Switching Threshold | VTH | (Notes 5, 8) | 0.7 3.4 | V |
| Switching Hysteresis | VHY | (Note 9) | 0.09 N/A | V |
| Output Low Voltage | VOL | At 4mA (Note 10) | 0.4 | V |
| Recovery Time (Note 2) | tREC | Standard speed, RPUP = 2.2k | 5 | μs |
|  |  | Overdrive speed, RPUP = 2.2k | 2 |  |
|  |  | Overdrive speed, direct ly prior to reset pulse; RPUP = 2.2k | 5 |  |
| Rising-Edge Hold-Off Time | tREH | (Note 11) | 0.6 2.0 | μs |
| Time Slot Duration (Note 2) | tSLOT | Standard speed | 65 | μs |
|  |  | Overdrive speed, VPUP > 4.5V | 8 |  |
|  |  | Overdrive speed (Note 12) | 9.5 |  |
| I/O PIN 1-Wire RESET, PRESENCE-DETECT CYCLE |  |  |  |  |
| Reset Low Time (Note 2) | tRSTL | Standard speed, VPUP > 4.5V | 480 720 | μs |
|  |  | Standard speed (Note 12) | 690 720 |  |
|  |  | Overdrive speed, VPUP > 4.5V | 48 80 |  |
|  |  | Overdrive speed (Note 12) | 70 80 |  |
| Presence-Detect High Time | tPDH | Standard speed, VPUP > 4.5V | 15 60 | μs |
|  |  | Standard speed (Note 12) | 15 63.5 |  |
|  |  | Overdrive speed (Note 12) | 2 7 |  |

DS1922E

ELECTRICAL CHARACTERISTICS (continued)

(VPUP = 3.0V to 5.25V.)

| W1L PARAMETER | SYMBOL | CONDITIONS | MIN TYP MAX | UNITS |
| --- | --- | --- | --- | --- |
| Presence-Detect Fall Time (Note 13) | tFPD | Standard speed, VPUP > 4.5V | 1.5 5 | μs |
|  |  | Standard speed | 1.5 8 |  |
|  |  | Overdrive speed | 0.15 1 |  |
| Presence-Detect Low Time | tPDL | Standard speed, VPUP > 4.5V | 60 240 | μs |
|  |  | Standard speed (Note 12) | 60 287 |  |
|  |  | Overdrive speed, VPUP > 4.5V (Note 12) | 7 24 |  |
|  |  | Overdrive speed (Note 12) | 7 28 |  |
| Presence-Detect Sample Time (Note 2) | tMSP | Standard speed, VPUP > 4.5V | 65 75 | μs |
|  |  | Standard speed | 71.5 75 |  |
|  |  | Overdrive speed | 8 9 |  |
| I/O PIN 1-Wire WRITE |  |  |  |  |
| Write-Zero Low Time (Notes 2, 14) | tW0L | Standard speed | 60 120 | μs |
|  |  | Overdrive speed, VPUP > 4.5V (Note 12) | 6 12 |  |
|  |  | Overdrive speed (Note 12) | 7.5 12 |  |
| Write-One Low Time (Notes 2, 14) | t | Standard speed | 5 15 | μs |
|  |  | Overdrive speed | 1 1.95 |  |
| I/O PIN 1-Wire READ |  |  |  |  |
| Read Low Time (Notes 2, 15) | tRL | Standard speed | 5 15 - | μs |
|  |  | Overdrive speed | 1 1.95 - |  |
| Read Sample Time (Notes 2, 15) | tMSR | Standard speed | tRL + 15 | μs |
|  |  | Overdrive speed | tRL + 1.95 |  |
| REAL-TIME CLOCK (RTC) |  |  |  |  |
| Accuracy |  |  | See the RTC Accuracy graph | Min/ Month |
| Frequency Deviat ion | F | 0°C to +125°C | -600 +60 | ppm |
| TEMPERATURE CONVERTER |  |  |  |  |
| Conversion Time (Note 16) | tCONV | 8-bit mode | 30 75 | ms |
|  |  | 16-bit mode (11 bit s) | 240 600 |  |
| Thermal Response Time Constant (Note 17) | RESP | iButton package | 130 | s |
| Conversion Error (Notes 18, 19) |  | +15°C to +110°C (Note 20) | ±7 | °C |
|  |  | +110°C to +140°C | -1.5 +1.5 |  |
| Temperature Cycles | NTCY | Cycle = ramp from +25°C to > +125°C and back to +25°C (Note 21) | 300 | Cycles |
| Operating Lifetime | tLIFE | Temperature > +125°C (Note 21) | 300 | Hours |

DS1922E

Note 1: 	Operation above +125°C is restricted  to mission operations  only. Communication and 1-Wire pin specifications are not specified for operation above +125°C.

Note 2: 	System requirement.

Note 3: 	Maximum allowable pullup resistance  is a function of the number of 1-Wire devices  in the system and 1-Wire recovery times. The specified value here applies to systems with only one device and with the minimum 1-Wire recovery times. For more heavily loaded systems, an active pullup such as that in the DS2480B can be required.

Note 4: 	Capacitance on the data pin could be 800pF when VPUP is first applied.  If a 2.2kΩ resistor is used to pull up the data line

2.5µs after VPUP has been applied,  the parasite capacitance does not affect normal communications.

Note 5: 	VTL and VTH are a function of the internal supply voltage, which is a function of VPUP and the 1-Wire recovery times. The

VTH and VTL maximum specifications are valid at VPUPMAX (5.25V). In any case, VTL < VTH < VPUP.

Note 6: 	Voltage below which, during a falling edge on I/O, a logic 0 is detected.

Note 7: 	The voltage on I/O must be less than or equal to VILMAX whenever the master drives the line low.

Note 8: 	Voltage above which, during a rising edge on I/O, a logic 1 is detected.

Note 9: 	After VTH is crossed  during a rising edge on I/O, the voltage on I/O must drop by VHY to be detected as logic 0.

Note 10:  The I-V characteristic is linear for voltages less than 1V.

Note 11:  The earliest recognition of a negative edge is possible  at tREH after VTH has been previously  reached.

Note 12:  Numbers in bold are not in compliance with the published iButton standards.  See the Comparison Table.

Note 13:  Interval during the negative edge on I/O at the beginning of a presence-detect pulse between the time at which the volt- age is 90% of VPUP and the time at which the voltage is 10% of VPUP.

Note 14:  ε in Figure 13 represents  the time required  for the pullup circuitry  to pull the voltage on I/O up from VIL to VTH. The actual

maximum duration for the master to pull the line low is tW1LMAX + tF - ε and tW0LMAX + tF - ε, respectively.

Note 15:  δ in Figure 13 represents  the time required  for the pullup circuitry  to pull the voltage on I/O up from VIL to the input high

threshold of the bus master. The actual maximum duration for the master to pull the line low is tRLMAX + tF.

Note 16:  To conserve battery power, use 8-bit temperature  logging  whenever possible.

Note 17:  This number was derived from a test conducted by Cemagref in Antony, France, in July 2000:

www.cemagref.fr/English/index.htm Test Report No. E42.

Note 18:  Includes  +0.1°C/-0.2°C calibration  chamber  measurement  uncertainty.

Note 19:  Warning: Not for use as the sole method of measuring  or tracking  temperature  in products and articles that could affect the health or safety of persons, plants, animals, or other living organisms,  including but not limited to foods, beverages, pharmaceuticals, medications, blood and blood products, organs, and flammable  and combustible products. User shall assure that redundant  (or other primary) methods of testing and determining the handling  methods, quality, and fitness of the articles and products should be implemented. Temperature  tracking  with this product,  where the health or safety of the aforementioned persons or things could be adversely affected,  is only recommended when supplemental or redun- dant information sources are used. Data-logger products are 100% tested and calibrated at time of manufacture by Maxim to ensure that they meet all data sheet parameters,  including temperature  accuracy. User shall be responsible for

proper use and storage of this product.  As with any sensor-based product,  user shall also be responsible for occasionally rechecking the temperature  accuracy of the product to ensure it is still operating  properly.

Note 20:  Guaranteed  by design and not production tested.

Note 21:  Devices leave the factory after having been run through a few cycles above +125°C. This is required  for calibration  of the device but should not affect lifetime of the device as specified. However, this process  results in a nonzero value in the Device Samples Counter register (0223h–0225h), which provides evidence the device has been factory calibrated.

COMPARISON TABLE

| PARAMETER | LEGACY VALUES |  |  |  | DS1922E VALUES |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | STANDARD SPEED (μs) |  | OVERDRIVE SPEED (μs) |  | STANDARD SPEED (μs) |  | OVERDRIVE SPEED (μs) |  |
|  | MIN | MAX | MIN | MAX | MIN | MAX | MIN | MAX |
| tSLOT (including tREC) | 61 | (undefined) | 7 | (undefined) | 65* | (undefined) | 9.5 | (undefined) |
| tRSTL | 480 | (undefined) | 48 | 80 | 690 | 720 | 70 | 80 |
| tPDH | 15 | 60 | 2 | 6 | 15 | 63.5 | 2 | 7 |
| tPDL | 60 | 240 | 8 | 24 | 60 | 287 | 7 | 28 |
| tW0L | 60 | 120 | 6 | 16 | 60 | 120 | 7.5 | 12 |

*Intentional change; longer recovery time requirement due to modified 1-Wire front-end.

Note: Numbers in bold are not in compliance with the published iButton standards.

iButton CAN PHYSICAL SPECIFICATION

| SIZE | See the Package Information section. |
| --- | --- |
| WEIGHT | Ca. 3.3 grams |
| SAFETY | Meets UL 913 (4th Edit.); Intrinsically Safe Apparatus, approval under Entity Concept for use in Class I, Division 1, Group A, B, C, and D Locations*. |

*Application pending.

RTC  Accuracy

RTC ACCURACY (TYPICAL)

2

0

-2

-4

DRIFT (MINUTES/MONTH)

-6

-8

-10

-12

-14

-16

-18

15	25

35	45	55	65	75	85

95	105

115

125

135

DS1922E

TEMPERATURE (°C)

Detailed Description

DS1922E

With its extended temperature range, the DS1922E is well suited to monitor processes that require tempera- tures well above the boiling point of water, such as pas- teurization of food items. Note that the initial sealing level of the DS1922E achieves the equivalent of IP56. Aging and use conditions can degrade  the integrity of the seal over time, so for applications  with significant exposure to liquids,  sprays, or other similar environ- ments, it is recommended  to place the DS1922E in the DS9107 iButton capsule. The DS9107 provides a water- tight enclosure that has been rated to IP68 (refer to Application  Note 4126: Understanding  the IP (Ingress Protection) Ratings of i Button Data Loggers  and Capsules). Software for setup and data retrieval through the 1-Wire interface is available for free download from the iButton website (www.ibutton.com). This software also includes drivers for the serial and USB port of a PC and routines to access the general-purpose memory for storing application- or equipment-specific data files.

Overview The block diagram in Figure 1 shows the relationships between the major control and memory sections of the DS1922E. The device has five main data components:

64-bit lasered ROM; 256-bit scratchpad;  576-byte gen- eral-purpose SRAM; two 256-bit register pages of time- keeping,  control, status, and counter registers, and passwords;  and 8192 bytes of data-logging  memory. Except for the ROM and the scratchpad,  all other mem- ory is arranged in a single linear address  space. The data-logging  memory, counter registers, and several other registers are read only for the user. Both register pages  are write protected  while the device  is pro- grammed for a mission. The password registers, one for a read password and another one for a read/write pass- word, can only be written, never read.

Figure 2 shows the hierarchical structure of the 1-Wire protocol.  The bus master must first provide one of the eight ROM function commands:  Read ROM, Match ROM, Search ROM, Conditional  Search ROM, Skip ROM, Overdrive-Skip ROM, Overdrive-Match  ROM, or Resume Command. Upon completion  of an Overdrive

ROM command byte executed at standard speed, the device enters Overdrive Mode, where all subsequent communication occurs at a higher speed. The protocol required  for  these  ROM function  commands  is described  in Figure 11. After a ROM function command is successfully executed, the memory and control func- tions become  accessible  and the master can provide any one of the eight available commands. The protocol for these memory and control function commands  is described  in Figure 9. All data is read and written least significant bit first.

Parasite Power The block diagram (Figure 1) shows the parasite-pow- ered circuitry. This circuitry “steals” power whenever the I/O input is high. I/O provides sufficient power as long as the specified  timing and voltage requirements are met. The advantages  of parasite power are two-fold: 1) By parasiting off this input, battery power is not consumed for 1-Wire ROM function commands, and 2) if the battery is exhausted for any reason, the ROM can still be read normally. The remaining circuitry of the DS1922E is sole- ly operated by battery energy.

64-Bit Lasered ROM Each DS1922E contains a unique ROM code that is 64 bits long. The first 8 bits are a 1-Wire family code. The next 48 bits are a unique serial number. The last 8 bits are a cyclic redundancy check (CRC) of the first 56 bits (see Figure 3 for details). The 1-Wire CRC is generated using a polynomial generator consisting of a shift regis- ter and XOR gates as shown in Figure 4. The polynomi- al is X8 + X5 + X4 + 1. Additional information about the

1-Wire CRC is available  in Application  Note 27: Understanding  and Using Cyclic  Redundancy  Checks with Maxim iButton Products.

The shift register bits are initialized to 0. Then, starting with the least significant bit of the family code, one bit at a time is shifted in. After the 8th bit of the family code has been entered, the serial number is entered. After the last bit of the serial number has been entered, the shift register contains the CRC value. Shifting in the 8 bits of CRC returns the shift register to all 0s.

1-Wire  PORT

I/O

ROM FUNCTION CONTROL

64-BIT LASERED ROM

PARASITE-POWERED CIRCUITRY

3V LITHIUM

MEMORY FUNCTION CONTROL

256-BIT SCRATCHPAD

DS1922E

GENERAL-PURPOSE SRAM

(512 BYTES)

32.768kHz

OSCILLATOR

INTERNAL TIMEKEEPING, CONTROL REGISTERS, AND COUNTERS

REGISTER PAGES (64 BYTES)

THERMAL	ADC SENSE

USER MEMORY (64 BYTES)

CONTROL LOGIC

DATA-LOG MEMORY

8KB

DS1922E

Figure 1. Block Diagram

BUS MASTER

1-Wire NET

OTHER DEVICES

COMMAND LEVEL:

DS1922E

AVAILABLE COMMANDS:	DATA FIELD AFFECTED:

1-Wire ROM FUNCTION COMMANDS

READ ROM MATCH ROM SEARCH ROM

CONDITIONAL SEARCH ROM SKIP ROM

RESUME

OVERDRIVE-SKIP ROM OVERDRIVE-MATCH ROM

64-BIT ROM, RC-FLAG

64-BIT ROM, RC-FLAG

64-BIT ROM, RC-FLAG

64-BIT ROM, RC-FLAG, ALARM FLAGS, SEARCH CONDITIONS RC-FLAG

RC-FLAG

RC-FLAG, OD-FLAG

64-BIT ROM, RC-FLAG, OD-FLAG

DS1922E-SPECIFIC MEMORY FUNCTION COMMANDS

WRITE SCRATCHPAD READ SCRATCHPAD

COPY SCRATCHPAD WITH PW READ MEMORY WITH PW AND CRC CLEAR MEMORY WITH PW

FORCED CONVERSION START MISSION WITH PW

STOP MISSION WITH PW

256-BIT SCRATCHPAD, FLAGS

256-BIT SCRATCHPAD

512-BYTE  DATA MEMORY,  REGISTERS, FLAGS, PASSWORDS MEMORY,  REGISTERS, PASSWORDS

MISSION TIMESTAMP, MISSION SAMPLES COUNTER, START DELAY, ALARM FLAGS, PASSWORDS

MEMORY ADDRESSES 020Ch TO 020Dh

FLAGS, TIMESTAMP, MEMORY ADDRESSES

020Ch TO 020Dh (WHEN LOGGING) FLAGS

Figure 2. Hierarchical Structure for 1-Wire Protocol

MSB

LSB

8-BIT

CRC CODE	48-BIT SERIAL NUMBER

8-BIT FAMILY CODE (41h)

MSB

LSB

MSB

LSB

MSB

LSB

Figure 3. 64-Bit Lasered ROM

POLYNOMIAL = X8 + X5 + X4 + 1

1ST STAGE

2ND STAGE

3RD STAGE

4TH STAGE

5TH STAGE

6TH STAGE

7TH STAGE

8TH STAGE

X0 	X1 	X2 	X3 	X4

X5 	X6 	X7 	X8

DS1922E

INPUT DATA

Figure 4. 1-Wire CRC Generator

Memory

Figure 5 shows the DS1922E memory map. Pages 0 to

15 contain 512 bytes of general-purpose  SRAM. The various registers to set up and control the device fill pages  16 and 17, called register pages  1 and 2 (see Figure 6 for details). Pages 18 and 19 can be used as extension of the general-purpose memory. The data-log logging  memory starts at address  1000h (page  128) and extends over 256 pages. The memory pages 20 to

127 are reserved for future extensions. The scratchpad is an additional page that acts as a buffer when writing to the SRAM memory or the register page.  The data

memory can be written at any time. The access type for the register pages is register-specific  and depends  on whether the device  is programmed  for a mission. Figure 6 shows the details. The data-log  memory is read only for the user. It is written solely under supervi- sion of the on-chip  control logic.  Due to the special behavior of the write access  logic  (write scratchpad, copy scratchpad),  it is recommended  to only write full pages at a time. This also applies to the register pages. See the Address Registers and Transfer Status section for details.

ADDRESS

0000h TO 001Fh

32-BYTE INTERMEDIATE  STORAGE SCRATCHPAD

PAGE 0

0020h TO 01FFh

PAGES 1 TO 15

0200h TO 021Fh

PAGE 16

0220h TO 023Fh

PAGE 17

0240h TO 025Fh

PAGE 18

0260h TO 027Fh

PAGE 19

0280h TO 0FFFh

PAGES 20 TO 127

1000h TO 2FFFh 32-BYTE GENERAL-PURPOSE SRAM (R/W)

GENERAL-PURPOSE SRAM (R/W)

32-BYTE REGISTER PAGE 1

32-BYTE REGISTER PAGE 2

GENERAL-PURPOSE SRAM (R/W)

GENERAL-PURPOSE SRAM (R/W)

(RESERVED FOR FUTURE EXTENSIONS)

DATA-LOG MEMORY (READ ONLY)

PAGES 128 TO 383

DS1922E

Figure 5. Memory Map

DS1922E

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 | FUNCTION | ACCESS* |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0200h | 0 | 10 Seconds |  |  | Single Seconds |  |  |  | Real- Time Clock Registers | R/W | R |
| 0201h | 0 | 10 Minutes |  |  | Single Minutes |  |  |  |  |  |  |
| 0202h | 0 | 12/24 | 20 Hour AM/PM | 10 Hour | Single Hours |  |  |  |  |  |  |
| 0203h | 0 | 0 | 10 Date |  | Single Date |  |  |  |  |  |  |
| 0204h | CENT | 0 | 0 | 10 Months | Single Months |  |  |  |  |  |  |
| 0205h | 10 Years |  |  |  | Single Years |  |  |  |  |  |  |
| 0206h | Low Byte |  |  |  |  |  |  |  | Sample Rate | R/W | R |
| 0207h | 0 | 0 | High Byte |  |  |  |  |  |  |  |  |
| 0208h | Low Threshold |  |  |  |  |  |  |  | Temperature Alarms | R/W | R |
| 0209h | High Threshold |  |  |  |  |  |  |  |  |  |  |
| 020Ah | (No Function with the DS1922E) |  |  |  |  |  |  |  | — | R/W | R |
| 020Bh | (No Function with the DS1922E) |  |  |  |  |  |  |  |  |  |  |
| 020Ch | Low Byte |  |  | 0 | 0 | 0 | 0 | 0 | Latest Temperature | R | R |
| 020Dh | High Byte |  |  |  |  |  |  |  |  |  |  |
| 020Eh | (No Function with the DS1922E) |  |  |  |  |  |  |  | — | R | R |
| 020Fh | (No Function with the DS1922E) |  |  |  |  |  |  |  |  |  |  |
| 0210h | 0 | 0 | 0 | 0 | 0 | 0 | ETHA | ETLA | Temperature Alarm Enable | R/W | R |
| 0211h | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | — | R/W | R |
| 0212h | 0 | 0 | 0 | 0 | 0 | 0 | EHSS | EOSC | RTC Control | R/W | R |
| 0213h | 1 | 1 | SUTA | RO | (X) | TLFS | 0 | ETL | Mis sion Control | R/W | R |
| 0214h | BOR | 1 | 1 | 1 | 0 | 0 | THF | TLF | Alarm Status | R | R |
| 0215h | 1 | 1 | 0 | WFTA | MEMCLR | 0 | MIP | 0 | General Status | R | R |
| 0216h | Low Byte |  |  |  |  |  |  |  | Start Delay Counter | R/W | R |
| 0217h | Center Byte |  |  |  |  |  |  |  |  |  |  |
| 0218h | High Byte |  |  |  |  |  |  |  |  |  |  |

*The left entry in the ACCESS column is valid between missions. The right entry shows the applicable access type while a mission is in progress.

Figure 6. DS1922E Register Pages Map

DS1922E

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 | FUNCTION | ACCESS* |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0219h | 0 | 10 Seconds |  |  | Single Seconds |  |  |  | Mis sion Timestamp | R | R |
| 021Ah | 0 | 10 Minutes |  |  | Single Minutes |  |  |  |  |  |  |
| 021Bh | 0 | 12/24 | 20 Hour AM/PM | 10 Hour | Single Hours |  |  |  |  |  |  |
| 021Ch | 0 | 0 | 10 Date |  | Single Date |  |  |  |  |  |  |
| 021Dh | CENT | 0 | 0 | 10 Months | Single Months |  |  |  |  |  |  |
| 021Eh | 10 Years |  |  |  | Single Years |  |  |  |  |  |  |
| 021Fh | (No Function; Reads 00h) |  |  |  |  |  |  |  | — | R | R |
| 0220h | Low Byte |  |  |  |  |  |  |  | Mis sion Samples Counter | R | R |
| 0221h | Center Byte |  |  |  |  |  |  |  |  |  |  |
| 0222h | High Byte |  |  |  |  |  |  |  |  |  |  |
| 0223h | Low Byte |  |  |  |  |  |  |  | Device Samples Counter | R | R |
| 0224h | Center Byte |  |  |  |  |  |  |  |  |  |  |
| 0225h | High Byte |  |  |  |  |  |  |  |  |  |  |
| 0226h | Conf iguration Code |  |  |  |  |  |  |  | Flavor | R | R |
| 0227h | EPW |  |  |  |  |  |  |  | PW Control | R/W | R |
| 0228h | First Byte |  |  |  |  |  |  |  | Read Access Password | W | — |
| … | … |  |  |  |  |  |  |  |  |  |  |
| 022Fh | Eighth Byte |  |  |  |  |  |  |  |  |  |  |
| 0230h | First Byte |  |  |  |  |  |  |  | Full Access Password | W | — |
| … | … |  |  |  |  |  |  |  |  |  |  |
| 0237h | Eighth Byte |  |  |  |  |  |  |  |  |  |  |
| 0238h | (No Function; All These Bytes Read 00h) |  |  |  |  |  |  |  | — | R | R |
| … |  |  |  |  |  |  |  |  |  |  |  |
| 023Fh |  |  |  |  |  |  |  |  |  |  |  |

*The left entry in the ACCESS column is valid between missions. The right entry shows the applicable access type while a mission is in progress.

Figure 6. DS1922E Register Pages Map (continued)

Detailed Register Descriptions

Timekeeping and Calendar The RTC and calendar  information is accessed  by reading/writing  the appropriate  bytes in the register page,  address  0200h to 0205h. For readings  to be valid, all RTC registers must be read sequentially start- ing at address 0200h. Some of the RTC bits are set to

0. These bits always read 0 regardless of how they are written. The number representation of the RTC registers is binary-coded decimal (BCD) format.

The DS1922E’s RTC can run in either 12hr or 24hr mode. Bit 6 of the Hours register (address 0202h) is defined as the 12hr or 24hr mode select bit. When high, the 12hr mode is selected. In the 12hr mode, bit 5 is the AM/PM bit with logic 1 being PM. In the 24hr mode, bit 5 is the 20hr bit (20hr to 23hr). The CENT bit, bit 7 of the Months register, can be written by the user. This bit changes  its state when the years counter transitions from 99 to 00.

The calendar logic is designed  to automatically com- pensate for leap years. For every year value that is either 00 or a multiple of 4, the device adds a 29th of February. This works correctly up to (but not including) the year 2100.

Sample Rate

The content of the Sample Rate register (addresses

0206h, 0207h) specifies the time elapse (in seconds if EHSS = 1, or minutes if EHSS = 0) between two tem- perature-logging  events. The sample rate can be any value from 1 to 16,383, coded  as an unsigned 14-bit binary number. If EHSS = 1, the shortest time between logging  events is 1s and the longest (sample rate =

3FFFh) is 4.55hr. If EHSS = 0, the shortest is 1min and the longest time is 273.05hr (sample rate = 3FFFh). The EHSS bit is located  in the RTC Control register at address  0212h. It is important that the user sets the EHSS bit accordingly  while setting the Sample Rate register. Writing a sample rate of 0000h results in a sample rate = 0001h, causing the DS1922E to log the temperature either every minute or every sec- ond depending upon the state of the EHSS bit.

RTC  Registers

DS1922E

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0200h | 0 | 10 Seconds |  |  | Single Seconds |  |  |  |
| 0201h | 0 | 10 Minutes |  |  | Single Minutes |  |  |  |
| 0202h | 0 | 12/24 | 20 Hour AM/PM | 10 Hour | Single Hours |  |  |  |
| 0203h | 0 | 0 | 10 Date |  | Single Date |  |  |  |
| 0204h | CENT | 0 | 0 | 10 Months | Single Months |  |  |  |
| 0205h | 10 Years |  |  |  | Single Years |  |  |  |

Note: During a mission, there is only read access to these registers. Bit cells marked “0” always read 0 and cannot be written to 1.

Sample Rate Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0206h | Sample Rate Low |  |  |  |  |  |  |  |
| 0207h | 0 | 0 | Sample Rate High |  |  |  |  |  |

Note: During a mission, there is only read access to these registers. Bit cells marked “0” always read 0 and cannot be written to 1.

Temperature Conversion The DS1922E’s temperature  range begins  at +15°C and ends at +140°C. Temperature values are repre- sented as an 8- or 16-bit unsigned binary number with a resolution of 0.5°C in 8-bit mode and 0.0625°C in

16-bit mode.

The higher temperature byte TRH is always valid. In

16-bit mode, only the three highest bits of the lower byte TRL are valid. The five lower bits all read 0. TRL is undefined if the device is in 8-bit temperature mode. An out-of-range temperature reading is indicated as 00h or

0000h when too cold and FFh or FFE0h when too hot.

With TRH and TRL representing the decimal equivalent of a temperature reading, the temperature value is cal- culated as:

ϑ(°C) = TRH/2 + 14 + TRL/512 (16-bit mode, TLFS = 1, see address 0213h)

ϑ(°C) = TRH/2 + 14 (8-bit mode, TLFS = 0, see address 0213h)

This equation is valid for converting temperature read- ings stored in the data-log memory as well as for data read from the Latest Temperature Conversion Result register.

To specify the temperature alarm thresholds, the previ- ous equations are resolved to:

TALM = 2 x ϑ(°C) - 28

Because the temperature alarm threshold is only one byte, the resolution or temperature increment is limited to 0.5°C. The TALM value must be converted into hexa- decimal  format before it can be written to one of the Temperature Alarm Threshold registers (Low Alarm address 0208h; High Alarm address 0209h). Independent of the conversion mode (8- or 16-bit), only the most significant byte of a temperature conversion is used to determine whether an alarm is generated.

DS1922E

Latest Temperature Conversion Result Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 | BYTE |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 020Ch | T2 | T1 | T0 | 0 | 0 | 0 | 0 | 0 | TRL |
| 020Dh | T10 | T9 | T8 | T7 | T6 | T5 | T4 | T3 | TRH |

Table 1. Temperature Conversion Examples

| MODE | TRH |  | TRL |  | (°C) |
| --- | --- | --- | --- | --- | --- |
|  | HEX | DECIMAL | HEX | DECIMAL |  |
| 8-Bit | 54h | 84 | — | — | 56.0 |
| 8-Bit | 17h | 23 | — | — | 25.5 |
| 16-Bit | 54h | 84 | 00h | 0 | 56.0000 |
| 16-Bit | 17h | 23 | 60h | 96 | 25.6875 |

Table 2. Temperature Alarm Threshold Examples

| (°C) | TALM |  |
| --- | --- | --- |
|  | HEX | DECIMAL |
| 65.5 | 67h | 103 |
| 30.0 | 20h | 32 |

Temperature Sensor Control Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0210h | 0 | 0 | 0 | 0 | 0 | 0 | ETHA | ETLA |

Note: During a mission, there is only read access to this register. Bits 2 to 7 have no function. They always read 0 and cannot be written to 1.

RTC  Control Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0212h | 0 | 0 | 0 | 0 | 0 | 0 | EHSS | EOSC |

Note: During a mission, there is only read access to this register. Bits 2 to 7 have no function. They always read 0 and cannot be written to 1.

Temperature Sensor Alarm The DS1922E has two Temperature Alarm Threshold registers (address 0208h, 0209h) to store values that determine whether a critical  temperature  has been reached.  A temperature  alarm is generated  if the device  measures an alarming temperature  and the alarm signaling is enabled. The bits ETLA and ETHA that enable the temperature alarm are located  in the Temperature Sensor Control register. The temperature alarm flags TLF and THF are found in the Alarm Status register at address 0214h.

DS1922E

Bit 1: Enable Temperature High Alarm (ETHA). This bit controls whether, during a mission, the temperature high alarm flag (THF) may be set, if a temperature con- version results in a value equal to or higher than the value in the Temperature High Alarm Threshold register. If ETHA is 1, temperature high alarms are enabled. If ETHA is 0, temperature high alarms are not generated.

Bit 0: Enable Temperature Low Alarm (ETLA). This bit controls whether, during a mission, the temperature low alarm flag (TLF) may be set, if a temperature con- version results in a value equal to or lower than the value in the Temperature Low Alarm Threshold register. If ETLA is 1, temperature low alarms are enabled. If ETLA is 0, temperature low alarms are not generated.

RTC Control To minimize the power consumption of a DS1922E, the RTC oscillator should be turned off when the device is not in use. The oscillator on/off bit is located in the RTC Control register. This register also includes  the EHSS bit, which determines whether the sample rate is speci- fied in seconds or minutes.

Bit 1: Enable High-Speed Sample (EHSS). This bit controls the speed of the sample rate counter. When set to logic 0, the sample rate is specified in minutes. When set to logic 1, the sample rate is specified in seconds.

Bit 0: Enable Oscillator (EOSC). This bit controls the crystal oscillator of the RTC. When set to logic 1, the oscillator starts. When written to logic 0, the oscillator stops and the device is in a low-power data-retention mode. This bit must be 1 for normal operation. A Forced Conversion or Start Mission command  automatically starts the RTC by changing the EOSC bit to logic 1.

Mission Control Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0213h | 1 | 1 | SUTA | RO | (X) | TLFS | 0 | ETL |

Note: During a mission, there is only read access to this register. Bits 6 and 7 have no function. They always read 1 and cannot be written to 0. Bits 1 and 3 control functions that are not available with the DS1922E. Bit 1 must be set to 0. Under this condition the setting of bit 3 becomes a “don’t care.”

Mission Control The DS1922E is set up for its operation by writing appropriate data to its special function registers, which are located  in the two register pages. The settings in the Mission Control register determine which format (8 or 16 bits) applies and whether old data can be over- written by new data once the data-log  memory is full. An additional control bit can be set to tell the DS1922E to wait with logging  data until a temperature alarm is encountered.

DS1922E

Bit 5: Start Mission Upon Temperature  Alarm (SUTA). This bit specifies  whether a mission begins immediately (includes delayed start) or if a temperature alarm is required to start the mission. If this bit is 1, the device performs an 8-bit temperature conversion at the selected sample rate and begins with data logging only if an alarming temperature (high alarm or low alarm) was found. The first logged  temperature is when the alarm occurred.  However, the mission sample counter does not increment. This functionality is guaranteed by design and not production tested.

Bit 4: Rollover Control (RO). This bit controls whether, during  a mission, the data-log  memory is overwritten with new data or whether data logging is stopped once the data-log memory is full. Setting this bit to 1 enables the rollover and data logging  continues at the begin- ning, overwriting previously collected  data. If this bit is

0, the logging  and conversions stop once the data-log memory is full. However, the RTC continues to run and the MIP bit remains set until the Stop Mission command is performed.

Bit 2: Temperature  Logging Format Selection (TLFS). This bit specifies the format used to store tem- perature readings in the data-log memory. If this bit is

0, the data is stored in 8-bit format. If this bit is 1, the

16-bit format is used (higher resolution). With 16-bit for- mat, the most significant  byte is stored at the lower address.

Bit 0: Enable Temperature Logging (ETL). To set up the device for a temperature-logging  mission, this bit must be set to logic 1. The recorded  temperature val- ues start at address 1000h.

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0214h | BOR | 1 | 1 | 1 | 0 | 0 | THF | TLF |

Note: There is only read access to this register. Bits 4 to 6 have no function. They always read 1. Bits 2 and 3 have no function with the DS1922E. They always read 0. The alarm status bits are cleared simultaneously when the Clear Memory Function is invoked. See memory and control functions for details.

General Status Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0215h | 1 | 1 | 0 | WFTA | MEMCLR | 0 | MIP | 0 |

Note: There is only read access to this register. Bits 0, 2, 5, 6, and 7 have no function.

Alarm Status The fastest way to determine whether a programmed temperature threshold was exceeded during a mission is through reading  the Alarm Status register. In a net- worked environment that contains multiple DS1922E iButtons, the devices  that encountered  an alarm can quickly  be identified  by means of the Conditional Search command  (see the 1-Wire ROM Function Commands  section).  The temperature  alarm only occurs if enabled  (see the Temperature Sensor Alarm section). The BOR alarm is always enabled.

DS1922E

Bit 7: Battery-On Reset Alarm (BOR). If this bit reads

1, the device  has performed  a power-on reset. This indicates that the device has experienced  a shock big enough to interrupt the internal battery power supply. The device may still appear functional, but it has lost its factory  calibration.  Any data found in the data-log memory should be disregarded.

Bit 1: Temperature High Alarm Flag (THF). If this bit reads 1, there was at least one temperature conversion during  a mission revealing a temperature equal to or higher than the value in the Temperature High Alarm register. A forced  conversion can affect the THF bit. This bit can also be set with the initial alarm in the SUTA = 1 mode.

Bit 0: Temperature Low Alarm Flag (TLF). If this bit reads 1, there was at least one temperature conversion during  a mission revealing a temperature equal to or lower than the value in the Temperature Low Alarm reg- ister. A forced  conversion can affect the TLF bit. This bit can also be set with the initial alarm in the SUTA = 1 mode.

General Status The information in the General Status register tells the host computer  whether a mission-related command was executed successfully.  Individual status bits indi- cate whether the DS1922E is performing  a mission, waiting for a temperature alarm to trigger the logging of data or whether the data from the latest mission has been cleared.

Bit 4: Waiting for Temperature Alarm (WFTA). If this bit reads 1, the Mission Start Upon Temperature Alarm was selected and the Start Mission command was suc- cessfully executed, but the device has not yet experi- enced the temperature alarm. This bit is cleared after a temperature  alarm event, but is not affected  by the Clear Memory command. Once set, WFTA remains set if a mission is stopped  before a temperature  alarm occurs. To clear WFTA manually before starting a new

mission, set the high temperature  alarm (address

0209h) to +15°C and perform a forced conversion.

Bit 3: Memory Cleared (MEMCLR). If this bit reads 1, the Mission Timestamp, mission samples counter, and all the alarm flags of the Alarm Status register have been cleared in preparation of a new mission. Executing the Clear Memory command  clears these memory sections. The MEMCLR bit returns to 0 as soon as a new mission is started by using the Start Mission command. The memory must be cleared for a mission to start.

Bit 1: Mission in Progress (MIP). If this bit reads 1, the device has been set up for a mission and this mis- sion is still in progress. The MIP bit returns from logic 1 to logic  0 when a mission is ended.  See the Start Mission and Stop Mission function commands.

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0216h | Delay Low Byte |  |  |  |  |  |  |  |
| 0217h | Delay Center Byte |  |  |  |  |  |  |  |
| 0218h | Delay H igh Byte |  |  |  |  |  |  |  |

Note: During a mission, there is only read access to this register.

Mission Timestamp Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0219h | 0 | 10 Seconds |  |  | Single Seconds |  |  |  |
| 021Ah | 0 | 10 Minutes |  |  | Single Minutes |  |  |  |
| 021Bh | 0 | 12/24 | 20 Hours AM/PM | 10 Hours | Single Hours |  |  |  |
| 021Ch | 0 | 0 | 10 Date |  | Single Date |  |  |  |
| 021Dh | CENT | 0 | 0 | 10 Months | Single Months |  |  |  |
| 021Eh | 10 Years |  |  |  | Single Years |  |  |  |

Note: There is only read access to this register.

Mission Samples Counter Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0220h | Low Byte |  |  |  |  |  |  |  |
| 0221h | Center Byte |  |  |  |  |  |  |  |
| 0222h | High Byte |  |  |  |  |  |  |  |

Note: There is only read access to this register.

Mission Start Delay The content of the Mission Start Delay Counter register tells how many minutes must expire from the time a mission was started until the first measurement of the mission takes place  (SUTA = 0) or until the device starts testing the temperature for a temperature alarm (SUTA = 1). The Mission Start Delay is stored as an unsigned 24-bit integer number. The maximum delay is

DS1922E

16,777,215min, equivalent to 11,650 days or roughly

31yr. If the start delay is nonzero and the SUTA bit is set to 1, first the delay must expire before the device starts testing for temperature alarms to begin logging data.

For a typical mission, the Mission Start Delay is 0. If a mission is too long for a single DS1922E to store all readings at the selected sample rate, one can use sev- eral devices and set the Mission Start Delay for the sec- ond device to start recording as soon as the memory of the first device  is full, and so on. The RO bit in the Mission Control register (address 0213h) must be set to

0 to prevent overwriting of collected  data once the data-log memory is full.

Mission Timestamp The Mission Timestamp register indicates the date and time of the first temperature  sample of the mission. There is only read access  to the Mission Timestamp register.

Mission Progress Indicator Depending  on settings in the Mission Control register (address 0213h), the DS1922E logs temperature in 8-bit or 16-bit format. The Mission Samples Counter together with the starting address and the logging format (8 or 16 bits) provide the information to identify valid blocks  of data that have been gathered  during  the current (MIP = 1) or latest mission (MIP = 0). See the Data-Log Memory Usage section for an illustration.

The number read from the mission samples counter indicates  how often the DS1922E woke up during  a mission to measure temperature. The number format is

24-bit unsigned integer. The mission samples counter is reset through the Clear Memory command.

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0223h | Low Byte |  |  |  |  |  |  |  |
| 0224h | Center Byte |  |  |  |  |  |  |  |
| 0225h | High Byte |  |  |  |  |  |  |  |

Note: There is only read access to this register.

Device Configuration Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 | PART |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0226h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | DS2422 |
|  | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | DS1923 |
|  | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | DS1922L |
|  | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | DS1922T |
|  | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | DS1922E |

Note: There is only read access to this register.

Password Control Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0227h | EPW |  |  |  |  |  |  |  |

Note: During a mission, there is only read access to this register.

Other Indicators The Device Samples Counter register is similar to the Mission Samples Counter register. During a mission this counter increments whenever the DS1922E wakes up to measure and log data and when the device is testing for a temperature alarm in SUTA mode. Between mis- sions the counter increments  whenever the Forced Conversion command is executed. This way the Device Samples Counter register functions like a gas gauge for the battery that powers the iButton.

DS1922E

The Device Samples Counter register is reset to zero when the iButton is assembled. The counter increments a couple of times during final test. The number format is

24-bit unsigned integer. The maximum number that can be represented in this format is 16,777,215.

The code  in the Device Configuration  register allows the master to distinguish between the DS2422 chip and different versions of the DS1922 iButtons. The Device Configuration Register table shows the codes assigned to the various devices.

Security by  Password The DS1922E is designed  to use two passwords  that control read access and full access.  Reading from or writing to the scratchpad  as well as the forced conver- sion command does not require a password. The pass- word  must be  transmitted  immediately  after the command  code  of the memory or control function. If password checking  is enabled, the password transmit- ted is compared to the passwords stored in the device. The data pattern stored in the Password Control regis- ter determines whether password checking is enabled.

To enable password  checking,  the EPW bits need to form a binary pattern of 10101010 (AAh). The default pattern of EPW is different from AAh. If the EPW pattern is different from AAh, any pattern is accepted  as long as it has a length of exactly 64 bits. Once enabled, changing  the passwords  and disabling  password checking  requires the knowledge  of the current full- access password.

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0228h | RP7 | RP6 | RP5 | RP4 | RP3 | RP2 | RP1 | RP0 |
| 0229h | RP15 | RP14 | RP13 | RP12 | RP11 | RP10 | RP9 | RP8 |
| … | … |  |  |  |  |  |  |  |
| 022Eh | RP55 | RP54 | RP53 | RP52 | RP51 | RP50 | RP49 | RP48 |
| 022Fh | RP63 | RP62 | RP61 | RP60 | RP59 | RP58 | RP57 | RP56 |

Note: There is only write access to this register. Attempting to read the password reports all zeros. The password cannot be changed while a mission is in progress.

Full Access Password Register

| ADDRESS | BIT 7 | BIT 6 | BIT 5 | BIT 4 | BIT 3 | BIT 2 | BIT 1 | BIT 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0230h | FP7 | FP6 | FP5 | FP4 | FP3 | FP2 | FP1 | FP0 |
| 0231h | FP15 | FP14 | FP13 | FP12 | FP11 | FP10 | FP9 | FP8 |
| … | … |  |  |  |  |  |  |  |
| 0236h | FP55 | FP54 | FP53 | FP52 | FP51 | FP50 | FP49 | FP48 |
| 0237h | FP63 | FP62 | FP61 | FP60 | FP59 | FP58 | FP57 | FP56 |

Note: There is only write access to this register. Attempting to read the password reports all zeros. The password cannot be changed while a mission is in progress.

Before enabling password checking, passwords for read-only access as well as for full access (read/write/control) must be written to the password registers. Setting up a password or enabling/dis- abling the password checking is done in the same way as writing data to a memory location; only the address is different. Since they are located in the same memory page, both passwords can be rede- fined at the same time.

DS1922E

The Read Access Password must be transmitted exact- ly in the sequence RP0, RP1…RP62, RP63. This pass- word only applies  to the Read Memory with CRC function. The DS1922E delivers the requested data only if the password transmitted by the master was correct or if password checking is not enabled.

The Full Access Password must be transmitted exactly in the sequence  FP0, FP1…FP62, FP63. It affects the functions Read Memory with CRC, Copy Scratchpad, Clear Memory, Start Mission, and Stop Mission. The DS1922E executes the command only if the password transmitted by the master was correct or if password checking is not enabled.

Due to the special behavior of the write-access  logic, the Password Control register and both passwords must be written at the same time. When setting up new passwords,  always verify (read back) the scratchpad before sending the Copy Scratchpad command. After a new password is successfully copied from the scratch- pad to its memory location, erase the scratchpad by fill- ing it with new data (Write Scratchpad  command). Otherwise, a copy  of the passwords  remains in the scratchpad for public read access.

ETL = 1

TLFS = 0

8192

8-BIT  ENTRIES

1000h

2FFFh

ETL = 1

TLFS = 1

4096

16-BIT  ENTRIES

1000h

WITH 16-BIT FORMAT, THE MOST SIGNIFICANT BYTE IS STORED AT THE LOWER ADDRESS.

2FFFh

alarm flags must be cleared  using the Memory Clear command. To enable the device for a mission, the ETL bit must be set to 1. These are general settings that must be made in any case, regardless  of the type of object to be monitored and the duration of the mission.

If alarm signaling is desired, the temperature alarm low and  high  thresholds  must  be  defined.  See the Temperature Conversion section for how to convert a temperature value into the binary code to be written to the threshold registers. In addition,  the temperature alarm must be enabled for the low and/or high thresh- old. This makes the device respond  to a Conditional

DS1922E

Figure 7. Temperature Logging

Data-Log Memory Usage

Once set up for a mission, the DS1922E logs the temper- ature measurements at equidistant time points entry after entry in its data-log memory. The data-log memory can store 8192 entries in 8-bit format or 4096 entries in 16-bit format (Figure 7). In 16-bit format, the higher 8 bits of an entry are stored at the lower address. Knowing the start- ing time point (Mission Timestamp) and the interval between temperature measurements, one can recon- struct the time and date of each measurement.

There are two alternatives to the way the DS1922E behaves after the data-log memory is filled with data. The user can program the device to either stop any further recording  (disable rollover) or overwrite the previously recorded data (enable rollover), one entry at a time, start- ing again at the beginning of the respective memory sec- tion. The contents of the Mission Samples Counter in conjunction  with the sample rate and the Mission Timestamp allow reconstructing the time points of all val- ues stored in the data-log memory. This gives the exact history over time for the most recent measurements taken. Earlier measurements cannot be reconstructed.

Missioning

The typical  task of the DS1922E iButton is recording temperature. Before the device can perform this func- tion, it needs to be set up properly. This procedure  is called missioning.

First, the DS1922E must have its RTC set to a valid time and date. This reference time can be the local time, or, when used inside of a mobile unit, UTC (also called GMT, Greenwich Mean Time), or any other time stan- dard that was agreed upon. The RTC oscillator must be running (EOSC = 1). The memory assigned to store the mission timestamp,  mission samples counter, and

Search command  (see the 1-Wire ROM Function Commands section), provided  that an alarming condi- tion has been encountered.

The setting of the RO bit (rollover enable) and sample rate depends  on the duration of the mission and the monitoring requirements. If the most recently logged data is important, the rollover should be enabled (RO =

1). Otherwise, one should estimate the duration of the mission in minutes and divide  the number by 8192 (8-bit format) or 4096 (16-bit format) to calculate the value of the sample rate (number of minutes between conversions). For example, if the estimated duration of a mission is 10 days (= 14400min), the 8192-byte capacity of the data-log memory would be sufficient to store a new 8-bit value every 1.8min (110s). If the DS1922E’s data-log  memory is not large enough to store all readings, one can use several devices and set the mission start delay to values that make the second device start logging as soon as the memory of the first device is full, and so on. The RO bit must be set to 0 to disable  rollover that would otherwise overwrite the logged data.

After the RO bit and the mission start delay are set, the sample rate must be written to the Sample Rate regis- ter. The sample rate can be any value from 1 to 16,383, coded  as an unsigned  14-bit binary  number. The fastest sample rate is one sample per second (EHSS =

1, sample rate = 0001h) and the slowest is one sample every 273.05hr (EHSS = 0, sample rate = 3FFFh). To get one sample every 6min, for example, the sample rate value must be set to 6 (EHSS = 0) or 360 decimal (equivalent to 0168h at EHSS = 1).

If there is a risk of unauthorized access to the DS1922E or manipulation of data, one should define passwords for read access and full access. Before the passwords become  effective, their use must be enabled. See the Security by Password section for more details.

The last step to begin  a mission is to issue the Start Mission command. As soon as it has received this com- mand, the DS1922E sets the MIP flag and clears the MEMCLR flag. With the Immediate/Delayed Start Mode (SUTA = 0), after as many minutes as specified  by the mission start delay are over, the device  wakes up, copies  the current  date  and  time to the Mission Timestamp register, and logs the first entry of the mis- sion. This increments both the mission samples counter and device samples counter. All subsequent log entries are made as specified by the value in the Sample Rate register and the EHSS bit.

If the Start Upon Temperature Alarm mode is chosen

(SUTA = 1) and temperature logging is enabled (ETL =

1), the DS1922E first waits until the start delay is over. Then the device wakes up in intervals as specified  by the sample rate and EHSS bit and measures the tem- perature. This increments the Device Samples Counter register only. The first sample of the mission is logged when the temperature alarm occurred.  However, the Mission Sample Counter does not increment. One sam- ple period later the Mission Timestamp register is set. From then on, both the Mission Samples Counter and Device Samples Counter registers increment at the same time. All subsequent  log entries are made as specified  by the value in the Sample Rate register and the EHSS bit.

The general-purpose memory operates independently of the other memory sections and is not write protected during a mission. All the DS1922E’s memory can be read

at any time, e.g., to watch the progress of a mission. Attempts to read the passwords read 00h bytes instead of the data that is stored in the password registers.

Memory Access

Address Registers and Transfer Status Because of the serial data transfer, the DS1922E employs three address registers called TA1, TA2, and E/S (Figure 8). Registers TA1 and TA2 must be loaded with the target address  to which the data is written or from which data is sent to the master upon a Read command.  Register E/S acts like a byte counter and transfer status register. It is used to verify data integrity with Write commands.  Therefore, the master only has read access to this register. The lower 5 bits of the E/S register indicate  the address  of the last byte that has been written to the scratchpad.  This address is called ending offset. The DS1922E requires that the ending offset is always 1Fh for a Copy Scratchpad to func- tion. Bit 5 of the E/S register, called PF or partial byte flag, is set if the number of data bits sent by the master is not an integer multiple of 8. Bit 6 is always a 0. Note that the lowest 5 bits of the target address also deter- mine the address within the scratchpad, where interme- diate storage of data begins.  This address  is called byte offset. If the target address for a Write command is

13Ch, for example, the scratchpad  stores incoming data beginning  at the byte offset 1Ch and is full after only 4 bytes. The corresponding  ending offset in this example  is 1Fh. For best  economy  of speed  and

DS1922E

BIT NUMBER 	7 	6 	5 	4 	3 	2 	1 	0

| TARGET ADDRESS (TA1) T7 | T6 | T5 | T4 | T3 | T2 | T1 | T0 |
| --- | --- | --- | --- | --- | --- | --- | --- |

| TARGET ADDRESS (TA2) T15 | T14 | T13 | T12 | T11 | T10 | T9 | T8 |
| --- | --- | --- | --- | --- | --- | --- | --- |

ENDING ADDRESS WITH DATA STATUS (E/S)

| (READ ONLY) AA | 0 | PF | E4 | E3 | E2 | E1 | E0 |
| --- | --- | --- | --- | --- | --- | --- | --- |

Figure 8. Address Registers

efficiency, the target address for writing should point to the beginning of a page, i.e., the byte offset is 0. Thus, the full 32-byte capacity of the scratchpad  is available, resulting also in the ending offset of 1Fh. The ending offset together with the PF flag are a means to support the master checking  the data integrity after a Write command.  The highest valued bit of the E/S register, called  authorization accepted  (AA), indicates  that a valid Copy command  for the scratchpad  has been received and executed. Writing data to the scratchpad clears this flag.

DS1922E

Writing with Verification To write data to the DS1922E, the scratchpad  must be used as intermediate storage. First, the master issues the Write Scratchpad  command to specify the desired target address, followed by the data to be written to the scratchpad.  In the next step, the master sends the Read Scratchpad  command  to read the scratchpad and to verify data integrity. As preamble to the scratch- pad  data, the DS1922E sends the requested  target address  TA1 and TA2 and the contents of the E/S Register. If the PF flag is set, data did not arrive cor- rectly in the scratchpad.  The master does not need to continue reading; it can start a new trial to write data to the scratchpad.  Similarly, a set AA flag indicates that the Write command was not recognized by the device. If everything went correctly, both flags are cleared and the ending offset indicates the address of the last byte written to the scratchpad.  Now the master can continue verifying every data bit. After the master has verified the data, it must send the Copy Scratchpad command. This command must be followed exactly by the data of the three address registers TA1, TA2, and E/S, as the master has read them verifying the scratchpad.  As soon as the DS1922E has received  these bytes, it copies the data to the requested location beginning at the target address.

Memory and Control

Function Commands

Figure 9 shows the protocols necessary for accessing the memory and the special  function registers of the DS1922E. An example on how to use these and other functions to set up the DS1922E for a mission is includ- ed in the Mission Example: Prepare and Start a New Mission section. The communication between the mas- ter and the DS1922E takes place  either at standard speed  (default, OD = 0) or at overdrive speed  (OD =

1). If not explicitly  set into the Overdrive Mode the DS1922E assumes standard  speed.  Internal memory access  during  a mission has priority over external access through the 1-Wire interface. This affects sever- al commands  in this section. See the Memory Access Conflicts section for details and solutions.

Write Scratchpad Command [0Fh] After issuing the Write Scratchpad command, the mas- ter must first provide  the 2-byte target address,  fol- lowed by the data to be written to the scratchpad.  The data is written to the scratchpad starting at the byte off- set T[4:0]. The master must send as many bytes as are needed to reach the ending offset of 1Fh. If a data byte is incomplete, its content is ignored and the partial byte flag PF is set.

When executing  the Write Scratchpad  command,  the CRC generator inside the DS1922E calculates  a CRC of the entire data stream, starting at the command code and ending  at the last data byte sent by the master (Figure 15). This CRC is generated  using the CRC-16 polynomial by first clearing the CRC generator and then shifting  in the command  code  (0Fh) of the Write Scratchpad  command,  the target addresses  TA1 and TA2 as supplied by the master, and all the data bytes. If the ending offset is 11111b, the master can send 16 read time slots and receive the inverted CRC-16 gener- ated by the DS1922E.

Note that both register pages are write protected  dur- ing a mission. Although the Write Scratchpad  com- mand works normally at any time, the subsequent copy scratchpad to a register page fails during a mission.

Read Scratchpad Command [AAh] This command  is used to verify scratchpad  data and target address.  After issuing the Read Scratchpad command, the master begins reading. The first 2 bytes are the target address. The next byte is the ending off- set/data status byte (E/S) followed by the scratchpad data beginning  at the byte offset T[4:0], as shown in Figure 8. The master can continue reading  data until the end of the scratchpad  after which it receives an

DS1922E

inverted  CRC-16 of  the  command  code,  target addresses TA1 and TA2, the E/S byte, and the scratch- pad data starting at the target address. After the CRC

is  read,  the  bus  master  reads  logic  1s  from  the

DS1922E until a reset pulse is issued.

Copy Scratchpad with Password [99h] This command is used to copy data from the scratch- pad to the writable memory sections. After issuing the Copy Scratchpad command, the master must provide a

3-byte authorization pattern, which can be obtained by reading  the scratchpad  for verification. This pattern must exactly match the data contained  in the three address  registers (TA1, TA2, E/S, in that order). Next, the master must transmit the 64-bit full-access  pass- word. If passwords  are enabled  and the transmitted password is different from the stored full-access pass- word, the Copy Scratchpad  with Password command fails. The device stops communicating  and waits for a reset pulse. If the password  was correct  or if pass- words were not enabled, the device tests the 3-byte authorization code.  If the authorization code  pattern matches, the AA flag is set and the copy begins. A pat- tern of alternating 1s and 0s is transmitted after the data has been copied  until the master issues a reset pulse. While the copy  is in progress,  any attempt to reset the part is ignored. Copy typically takes 2µs per byte.

The data to be copied  is determined  by the three address registers. The scratchpad data from the begin- ning offset through the ending offset are copied,  start- ing at the target address. The AA flag remains at logic

1 until it is cleared by the next Write Scratchpad  com- mand. With suitable password,  the copy  scratchpad always functions for the 16 pages of data memory and the 2 pages of calibration memory. While a mission is in

progress, write attempts to the register pages  are not successful. The AA bit remaining at 0 indicates this.

Read Memory with Password and CRC [69h]

The Read Memory with CRC command  is the general

function to read from the device. This command gener- ates and transmits a 16-bit CRC following the last data byte of a memory page.

After having sent the command  code  of the Read

Memory with CRC command, the bus master sends a

2-byte address  that indicates a starting byte location. Next, the master must transmit one of the 64-bit pass- words. If passwords  are enabled  and the transmitted password does not match one of the stored passwords, the Read Memory with Password and CRC command fails. The device stops communicating  and waits for a reset pulse. If the password  was correct  or if pass- words were not enabled, the master reads data from the DS1922E beginning  from the starting address and continuing until the end of a 32-byte page is reached. At that point the bus master sends 16 additional read data time slots and receives the inverted 16-bit CRC. With subsequent  read data time slots the master receives data starting at the beginning  of the next memory page followed again by the CRC for that page. This sequence continues until the bus master resets the device. When trying to read the passwords or memory areas that are marked as “reserved,”  the DS1922E transmits 00h or FFh bytes, respectively.  The CRC at the end of a 32-byte memory page  is based  on the data as it was transmitted.

With the initial pass through the Read Memory with CRC flow, the 16-bit CRC value is the result of shifting the command byte into the cleared CRC generator fol- lowed by the two address bytes and the contents of the data memory. Subsequent passes through the Read Memory with CRC flow generate a 16-bit CRC that is the result of clearing the CRC generator and then shift- ing in the contents of the data memory page. After the

16-bit CRC of the last page  is read, the bus master receives logic 1s from the DS1922E until a reset pulse

is issued. The Read Memory with CRC command sequence can be ended at any point by issuing a reset pulse.

MASTER Tx MEMORY OR CONTROL FUNCTION COMMAND

FROM ROM FUNCTIONS FLOWCHART (FIGURE 11)

0Fh	N WRITE SCRATCHPAD?

AAh	N READ SCRATCHPAD?

99h

COPY SCRATCHPAD 	N

[WITH PW]

TO FIGURE 9b

Y	Y	Y

MASTER Tx

TA1 [T7:T0]

MASTER Tx

TA2 [T15:T8]

MASTER Rx

TA1 [T7:T0]

MASTER Rx

TA2 [T15:T8]

MASTER Tx

TA1 [T7:T0], TA2 [T15:T8]

MASTER Tx

E/S BYTE

AUTHORIZATION CODE

DS1922E SETS SCRATCHPAD OFFSET = [T4:T0]  AND CLEARS (PF, AA)

MASTER Rx ENDING OFFSET WITH DATA STATUS (E/S)

MASTER Tx

64 BITS [PASSWORD]

MASTER Tx DATA BYTE TO SCRATCHPAD  OFFSET

DS1922E SETS SCRATCHPAD  OFFSET = [T4:T0]

PASSWORD	N ACCEPTED?

DS1922E INCREMENTS SCRATCHPAD OFFSET

DS1922E SETS [E4:E0] = SCRATCHPAD OFFSET

DS1922E INCREMENTS SCRATCHPAD OFFSET

MASTER Rx DATA BYTE FROM SCRATCHPAD OFFSET

Y

AUTHORIZATION	N CODE MATCH?

Y MASTER Tx RESET?

N

Y MASTER Tx RESET?

N

Y AA = 1

N 		SCRATCHPAD OFFSET = 11111b?

N 		SCRATCHPAD OFFSET = 11111b?

DS1922E COPIES SCRATCHPAD DATA TO MEMORY

Y		PARTIAL	Y BYTE WRITTEN?

Y 	MASTER Rx "1"s	MASTER Rx "1"s

Y		N MASTER Tx RESET?

N

MASTER Rx CRC-16  OF COMMAND,  ADDRESS DATA

PF = 1

MASTER Rx CRC-16  OF COMMAND,  ADDRESS DATA, E/S BYTE, AND DATA STARTING AT THE TARGET ADDRESS

Y

MASTER  Tx RESET?

N 	COPYING FINISHED

Y DS1922E Tx "0"

N MASTER Tx RESET?

Y

Y

MASTER Tx RESET?

N MASTER Rx "1"s

Y MASTER Tx RESET?

N

DS1922E

N 	DS1922E Tx "1"

MASTER Rx "1"s

N

MASTER  Tx RESET?

Y

FROM FIGURE 9b

TO ROM FUNCTIONS FLOWCHART (FIGURE 11)

Figure 9a. Memory/Control Function Flowchart

FROM FIGURE 9a

69h

READ MEMORY [WITH	N PW] AND CRC

96h

CLEAR MEMORY	N [WITH PW]

55h	N FORCED CONVERSION?

TO FIGURE 9c

Y	Y	Y

MASTER Tx

TA1 [T7:T0], TA2 [T15:T8]

MASTER Tx

64 BITS [PASSWORD]

MASTER Tx

FFh DUMMY BYTE

DECISION MADE BY DS1922E

MASTER Tx

64 BITS [PASSWORD]

MASTER Tx

FFh DUMMY BYTE

MISSION IN	Y PROGRESS?

DECISION MADE BY MASTER

PASSWORD	N ACCEPTED?

Y DS1922E SETS

MEMORY ADDRESS = [T15:T0]

PASSWORD	N ACCEPTED?

Y

MISSION IN	Y PROGRESS?

N

DS1922E PERFORMS A TEMPERATURE CONVERSION

DS1922E  COPIES RESULT TO ADDRESS 020C/Dh

MASTER Rx DATA BYTE FROM MEMORY ADDRESS

Y

MASTER  Tx RESET?

N

DS1922E INCREMENTS ADDRESS COUNTER

N DS1922E CLEARS

MISSION TIMESTAMP,

MISSION SAMPLES COUNTER, ALARM FLAGS

DS1922E SETS MEMCLR = 1

N

MASTER  Tx RESET?

Y

N END OF PAGE?

Y

MASTER Rx CRC-16  OF COMMAND,  ADDRESS, DATA (1ST PASS); CRC-16 OF DATA (SUBSEQUENT PASSES)

N

MASTER  Tx RESET?

Y

N MASTER Tx RESET

CRC OK?

Y

N END OF MEMORY?

Y

MASTER Rx "1"s

N MASTER Tx RESET?

Y

TO FIGURE 9a

FROM FIGURE 9c

DS1922E

Figure 9b. Memory/Control Function Flowchart (continued)

FROM FIGURE 9b

CCh

START MISSION	N [WITH PW]

Y

MISSION START DELAY PROCESS

33h

STOP MISSION	N [WITH PW]

Y

MASTER Tx

64 BITS [PASSWORD]

MASTER Tx

FFh DUMMY BYTE

START DELAY	Y COUNTER = 0?

N

MASTER Tx

64 BITS [PASSWORD]

MASTER Tx

FFh DUMMY BYTE

DS1922E WAITS FOR 1 MINUTE

N 	PASSWORD ACCEPTED?

DS1922E DECREMENTS START DELAY COUNTER

PASSWORD	N ACCEPTED?

Y	Y

Y	MISSION IN PROGRESS?

N SUTA = 1?

MISSION IN	N PROGRESS?

N 	Y	Y

N

MEMCLR = 1?

Y

DS1922E SETS WFTA = 1

DS1922E  WAITS ONE SAMPLE PERIOD

DS1922E SETS MIP = 0, WFTA = 0

DS1922E SETS MIP = 1, MEMCLR = 0

Y

MIP = 0?

N MASTER Tx RESET?

DS1922E INITIATES MISSION START DELAY PROCESS

N 	Y DS1922E PERFORMS 8-BIT

TEMPERATURE CONVERSION

N

MASTER  Tx RESET?

Y

TEMPERATURE	N ALARM?

Y

DS1922E

DS1922E SETS WFTA = 0

DS1922E  WAITS ONE SAMPLE PERIOD

DS1922E  COPIES RTC DATA TO MISSION TIMESTAMP REGISTER

DS1922E STARTS LOGGING TAKING FIRST SAMPLE

END OF PROCESS

TO FIGURE 9b

Figure 9c. Memory/Control Function Flowchart (continued)

Clear Memory with Password [96h] The Clear Memory with Password command is used to prepare the device for another mission. This command is only executed if no mission is in progress. After the command code the master must transmit the 64-bit full- access password  followed by an FFh dummy byte. If passwords are enabled and the transmitted password is different from the stored full-access  password  or a mission is in progress, the Clear Memory with Password command fails. The device stops communi- cating and waits for a reset pulse. If the password was correct or if passwords  were not enabled, the device clears  the Mission Timestamp  register,  Mission Samples Counter register, and all alarm flags of the Alarm Status register. After these cells are cleared, the MEMCLR bit of the General Status register reads 1 to indicate the successful execution of the Clear Memory with Password command.  Clearing of the data-log memory is not necessary because the mission samples counter indicates  how many entries in the data-log memory are valid.

DS1922E

Forced Conversion [55h] The Forced Conversion command can be used to mea- sure the temperature without starting a mission. After the command  code,  the master must send one FFh byte to get the conversion  started. The conversion result is found as a 16-bit value in the Latest Temperature Conversion Result register. This com- mand is only executed if no mission is in progress (MIP

= 0). It cannot be interrupted  and takes maximum

600ms to complete.  During this time memory access through the 1-Wire interface is blocked.  The device behaves the same way as during  a mission when the sampling  interferes with a memory/control function command.  See the Memory Access  Conflicts section for details.

Start Mission with Password [CCh] The DS1922E uses a control function command to start a mission. A new mission can only be started if the pre- vious mission has been ended  and the memory has been cleared.  After the command  code,  the master must transmit the 64-bit full-access password followed by an FFh dummy byte. If passwords are enabled and the transmitted password  is different from the stored full-access  password  or a mission is in progress, the Start Mission with Password command fails. The device stops communicating and waits for a reset pulse. If the password  was correct  or if passwords  were not enabled, the device starts a mission. If SUTA = 0, the sampling begins as soon as the mission start delay is over. If SUTA = 1, the first sample is written to the data-

log memory at the time the temperature alarm occurred.  However, the mission sample counter does not increment. One sample period  later, the Mission Timestamp register is set and the regular sampling and logging  begins. While the device is waiting for a tem- perature alarm to occur, the WFTA flag in the General Status register reads 1. During a mission there is only read access to the register pages.

Stop Mission with Password [33h] The DS1922E uses a control function command to stop a mission. Only a mission that is in progress can be stopped.  After the command  code,  the master must transmit the 64-bit full-access password followed by a FFh dummy byte. If passwords  are enabled  and the transmitted password  is different from the stored full- access password  or a mission is not in progress, the Stop Mission with Password command fails. The device stops communicating and waits for a reset pulse. If the password  was correct  or if passwords  were not enabled, the device clears the MIP bit in the General Status register and restores write access to the register pages. The WFTA bit is not cleared.  See the descrip- tion of the General Status register for a method to clear the WFTA bit.

Memory Access Conflicts

While a mission is in progress  or while the device is waiting for a temperature alarm to start a mission, peri- odically  a temperature sample is taken and logged. This “internal activity” has priority over 1-Wire communi- cation. As a consequence,  device-specific  commands (excluding ROM function commands and 1-Wire reset) do not perform properly  when internal and “external” activities interfere with each other. Not affected are the commands  Start Mission, Forced Conversion, and Clear Memory, because they are not applicable  while a

mission is in progress or while the device is waiting for a temperature alarm. Table 3 explains how the remain- ing five commands are affected by internal activity, how to detect this interference, and how to work around it.

The interference is more likely to be seen with a high- sample rate (one sample every second) and with high- resolution logging,  which can last up to 600ms. With lower sample rates, interference may hardly be visible at all. In any case, when writing driver software it is important to know about the possibility of interference and to take measures to work around it.

DS1922E

Table 3. Memory Access Conflicts and Solutions

| COMMAND | INDICATION OF INTERFERENCE | SOLUTION |
| --- | --- | --- |
| Write Scratchpad | The CRC-16 at the end of the command flow reads FFFFh. | Wait 0.5s, 1-Wire reset, address the device, repeat Write Scratchpad with the same data, and check the validit y of the CRC-16 at the end of the command flow. Alternatively, use Read Scratchpad to verify data integrit y. |
| Read Scratchpad | The data read changes to FFh bytes or all bytes received are FFh, including the CRC at the end of the command flow. | Wait 0.5s, 1-Wire reset, address the device, repeat Read Scratchpad, and check the valid it y of the CRC-16 at the end of the command flow. |
| Copy Scratchpad | The device behaves as if the authorizat ion code or password was not valid, or as if the copy function would not end. | Wait 0.5s, 1-Wire reset, address the device, issue Read Scratchpad, and check the AA bit of the E/S byte. If the AA bit is set, Copy Scratchpad was successful. |
| Read Memory with CRC | The data read changes to all FFh bytes or all bytes received are FFh, including the CRC at the end of the command flow, despite a valid password. | Wait 0.5s, 1-Wire reset, address the device, repeat Read Memory with CRC, and check the validit y of the CRC-16 at the end of the memory page. |
| Stop Miss ion | The General Status register at address 0215h reads FFh or the MIP bit is 1 while bit s 0, 2, and 5 are 0. | Wait 0.5s, 1-Wire reset, address the device, and repeat Stop Mission. Perform a 1-Wire reset, address the device, read the General Status register at address 0215h, and check the MIP bit. If the MIP bit is 0, Stop Mission was successful. |

1-Wire Bus System

The 1-Wire bus is a system that has a single bus mas- ter and  one or more slaves. In all instances  the DS1922E is a slave device. The bus master is typically a microcontroller. The discussion of this bus system is broken down into three topics: hardware configuration, transaction sequence,  and 1-Wire signaling  (signal types and timing). The 1-Wire protocol  defines bus transactions in terms of the bus state during  specific time slots that are initiated on the falling edge of sync pulses from the bus master.

Hardware Configuration

The 1-Wire bus has only a single line by definition; it is important that each device on the bus be able to drive it at the appropriate time. To facilitate this, each device attached  to the 1-Wire bus must have open-drain  or three-state outputs. The 1-Wire port of the DS1922E is open drain with an internal circuit  equivalent to that shown in Figure 10.

A multidrop bus consists of a 1-Wire bus with multiple slaves attached. At standard speed the 1-Wire bus has a maximum data rate of 16.3kbps. The speed can be boosted to 142kbps by activating the Overdrive Mode. The DS1922E is not guaranteed to be fully compliant to the iButton standard. Its maximum data rate in standard speed  is 15.4kbps  and 125kbps in overdrive speed.

The value of the pullup  resistor primarily depends  on the network size and load conditions.  The DS1922E requires a pullup  resistor of maximum 2.2kΩ at any speed.

The idle state for the 1-Wire bus is high. If for any rea- son a transaction needs to be suspended,  the bus must be left in the idle state if the transaction is to resume. If this does not occur and the bus is left low for more than 16µs (overdrive speed) or more than 120µs (standard speed), one or more devices on the bus may be reset. Note that the DS1922E does not quite meet the full 16µs maximum low time of the normal 1-Wire bus overdrive timing. With the DS1922E the bus must be left low for no longer than 12µs at overdrive to ensure that no DS1922E on the 1-Wire bus performs a reset. The DS1922E communicates properly when used in conjunction with a DS2480B or DS2490 1-Wire driver and adapters that are based on these driver chips.

Transaction Sequence

The protocol  for accessing  the DS1922E through the

1-Wire port is as follows:

• Initialization

• ROM Function Command

• Memory/Control Function Command

• Transaction/Data

VPUP

BUS MASTER

Rx

Tx

OPEN-DRAIN

RPUP

Rx = RECEIVE Tx = TRANSMIT

DATA

IL

DS1922E 1-Wire PORT

Rx

Tx

PORT PIN 	100Ω MOSFET

DS1922E

Figure 10. Hardware Configuration

Initialization

DS1922E

All transactions on the 1-Wire bus begin with an initial- ization sequence.  The initialization sequence  consists of a reset pulse transmitted by the bus master followed by presence pulse(s) transmitted by the slave(s). The presence  pulse lets the bus master know that the DS1922E is on the bus and is ready to operate. For more details, see the 1-Wire Signaling section.

1-Wire ROM  Function Commands

Once the bus master has detected  a presence, it can issue one of the eight ROM function commands that the DS1922E supports. All ROM function commands are 8 bits long. A list of these commands  follows (see the flowchart in Figure 11).

Read ROM  [33h] This command  allows the bus master to read the DS1922E’s 8-bit family code, unique 48-bit serial number, and 8-bit CRC. This command can only be used if there is a single slave on the bus. If more than one slave is pre- sent on the bus, a data collision occurs when all slaves try to transmit at the same time (open drain produces  a wired-AND result). The resultant family code and 48-bit serial number results in a mismatch of the CRC.

Match ROM  [55h] The Match ROM command, followed by a 64-bit ROM sequence, allows the bus master to address a specific DS1922E on a multidrop bus. Only the DS1922E that exactly matches the 64-bit ROM sequence responds to the following memory function command. All other slaves wait for a reset pulse. This command can be used with a single device or multiple devices on the bus.

Search ROM  [F0h] When a system is initially brought  up, the bus master might not know the number of devices  on the 1-Wire bus or their registration numbers. By taking advantage of the wired-AND property  of the bus, the master can use a process of elimination to identify the registration

numbers of all slave devices. For each bit of the regis- tration number, starting with the least significant bit, the bus master issues a triplet of time slots. On the first slot, each slave device participating  in the search outputs the true value of its registration number bit. On the sec- ond slot, each slave device participating  in the search outputs the complemented value of its registration num- ber bit. On the third slot, the master writes the true value of the bit to be selected. All slave devices that do not match the bit written by the master stop participat- ing in the search. If both of the read bits are zero, the master knows that slave devices exist with both states of the bit. By choosing  which state to write, the bus master branches in the ROM code tree. After one com- plete pass, the bus master knows the registration num- ber of a single device.  Additional  passes identify the registration numbers of the remaining devices. Refer to Application  Note 187: 1-Wire Search Algorithm for a detailed discussion, including an example.

Conditional Search ROM  [ECh] The Conditional Search ROM command operates simi- larly to the Search ROM command  except  that only those devices that fulfill certain conditions participate in the search. This function provides an efficient means for the bus master to identify devices  on a multidrop system that have to signal an important event. After each pass of the conditional  search that successfully determined the 64-bit ROM code for a specific  device on the multidrop bus, that particular device can be indi- vidually accessed as if a Match ROM had been issued, since all other devices have dropped  out of the search process and are waiting for a reset pulse.

The DS1922E responds to the Conditional Search ROM command  if one of the three alarm flags of the Alarm Status register (address 0214h) reads 1. The tempera- ture alarm only occurs if enabled (see the Temperature Sensor Alarm section).  The BOR alarm is always enabled. The first alarm that occurs makes the device respond to the Conditional Search ROM command.

Skip ROM  [CCh] This command can save time in a single-drop bus sys- tem by allowing the bus master to access the memory functions without providing  the 64-bit ROM code. For example, if more than one slave is present on the bus and a Read command is issued following the Skip ROM command, data collision occurs on the bus as multiple slaves transmit simultaneously (open-drain  pulldowns produce a wired-AND result).

DS1922E

Resume [A5h] The DS1922E must be accessed several times before a mission starts. In a multidrop environment this means that the 64-bit ROM code after a Match ROM command must be repeated for every access.  To maximize the data throughput in a multidrop environment, the Resume function was implemented. This function checks the sta- tus of the RC bit and, if it is set, directly transfers control to the memory/control functions, similar to a Skip ROM command.  The only way to set the RC bit is through successfully executing the Match ROM, Search ROM, or Overdrive-Match  ROM command.  Once the RC bit is set, the device can repeatedly be accessed through the Resume command function. Accessing  another device on the bus clears the RC bit, preventing  two or more devices from simultaneously responding to the Resume command function.

Overdrive-Skip ROM  [3Ch] On a single-drop  bus this command can save time by allowing the bus master to access the memory/control functions  without providing  the 64-bit ROM code. Unlike the normal Skip ROM command, the Overdrive-

Skip ROM command sets the DS1922E in the Overdrive Mode (OD = 1). All communication following this com- mand must occur at overdrive speed until a reset pulse of minimum 690µs duration resets all devices  on the bus to standard speed (OD = 0).

When issued on a multidrop bus, this command sets all overdrive-supporting  devices  into Overdrive Mode. To subsequently address  a specific  overdrive-supporting device,  a reset pulse at overdrive speed  must be issued followed by a Match ROM or Search ROM com- mand  sequence.  This speeds  up the time for the search process.  If more than one slave supporting overdrive is present on the bus and the Overdrive-Skip ROM command is followed by a Read command, data collision occurs on the bus as multiple slaves transmit simultaneously (open-drain pulldowns produce a wired- AND result).

Overdrive-Match ROM  [69h]

The Overdrive-Match  ROM command  followed by a

64-bit ROM sequence  transmitted at overdrive speed allows the bus master to address  a specific  DS1922E on a multidrop  bus and to simultaneously set it in Overdrive Mode. Only the DS1922E that exactly match- es the 64-bit ROM sequence  responds  to the subse- quent memory/control function command. Slaves already in Overdrive Mode from a previous Overdrive- Skip ROM or successful  Overdrive-Match  ROM com- mand remain in Overdrive Mode. All overdrive-capable slaves return to standard speed at the next reset pulse of minimum 690µs duration. The Overdrive-Match ROM command  can  be  used  with a single  or multiple devices on the bus.

FROM MEMORY FUNCTIONS FLOWCHART (FIGURE 9)

BUS MASTER Tx

RESET PULSE

FROM FIGURE 11b

OD	N RESET PULSE?

OD = 0

Y

BUS MASTER Tx ROM FUNCTION COMMAND

DS1922E Tx

PRESENCE PULSE

33h

READ ROM	N COMMAND?

55h

MATCH ROM	N COMMAND?

F0h SEARCH ROM COMMAND?

ECh

N 	CONDITIONAL SEARCH      N

COMMAND?

TO FIGURE 11b

Y	Y	Y	Y

RC = 0

| RC = 0 RC = 0 RC = 0 DS1922E Tx BIT 0 |  |
| --- | --- |
| DS1922E Tx BIT 0 |  |
| MASTER Tx BIT 0 |  |

N 	CONDITION MET?

| Y DS1922E Tx BIT 0 |  |
| --- | --- |
| DS1922E Tx BIT 0 |  |
| MASTER Tx BIT 0 |  |

DS1922E Tx FAMILY CODE (1 BYTE)

MASTER Tx BIT 0

BIT 0 MATCH?       N	N

| BIT 0 MATCH? DS1922E Tx BIT 1 |  |
| --- | --- |
| DS1922E Tx BIT 1 |  |
| MASTER Tx BIT 1 |  |

N       BIT 0 MATCH?

DS1922E Tx SERIAL NUMBER (6 BYTES)

Y	Y Y

MASTER Tx BIT 1

| BIT 1 MATCH? N N DS1922E Tx BIT 1 |  |
| --- | --- |
| DS1922E Tx BIT 1 |  |
| MASTER Tx BIT 1 |  |

BIT 1 MATCH?

N       BIT 1 MATCH?

| Y Y Y DS1922E Tx BIT 63 |  |
| --- | --- |
| DS1922E Tx BIT 63 |  |
| MASTER Tx BIT 63 |  |

DS1922E Tx

CRC BYTE

MASTER Tx BIT 63

BIT 63 MATCH?      N	N

| BIT 63 MATCH? DS1922E Tx BIT 63 |  |
| --- | --- |
| DS1922E Tx BIT 63 |  |
| MASTER Tx BIT 63 |  |

N      BIT 63 MATCH?

Y	Y	Y

RC = 1

RC = 1

RC = 1

TO FIGURE 11b

FROM FIGURE 11b

DS1922E

TO MEMORY FUNCTIONS FLOWCHART (FIGURE 9)

Figure 11a. ROM Functions Flowchart

TO FIGURE 11a

FROM FIGURE 11a

CCh

SKIP ROM	N COMMAND?

A5h

RESUME	N COMMAND?

3Ch

OVERDRIVE-	N SKIP ROM?

69h

OVERDRIVE-	N MATCH ROM?

Y	Y	Y	Y

RC = 0

N RC = 1?

RC = 0; OD = 1	RC = 0; OD = 1

Y

MASTER Tx	Y RESET?

MASTER Tx BIT 0

N

MASTER Tx	Y RESET?

BIT 0 MATCH?        N

(SEE NOTE) OD = 0

N

Y

MASTER Tx BIT 1

BIT 1 MATCH?        N

(SEE NOTE) OD = 0

Y

MASTER Tx BIT 63

BIT 63 MATCH?       N

(SEE NOTE) OD = 0

Y

FROM FIGURE 11a

TO FIGURE 11a

RC = 1

DS1922E

NOTE:  THE OD FLAG REMAINS  AT 1 IF THE DEVICE WAS ALREADY AT OVERDRIVE SPEED BEFORE THE OVERDRIVE-MATCH  ROM COMMAND WAS ISSUED.

Figure 11b. ROM Functions Flowchart (continued)

1-Wire Signaling

The DS1922E requires strict protocols to ensure data integrity. The protocol consists of four types of signaling on one line: reset sequence with reset pulse and pres- ence pulse, write-zero, write-one, and read-data. Except for the presence pulse, the bus master initiates all these signals. The DS1922E can communicate at two different speeds:  standard speed  and overdrive speed. If not explicitly  set into the Overdrive Mode, the DS1922E communicates at standard speed. While in Overdrive Mode the fast timing applies to all waveforms.

To get from idle to active, the voltage on the 1-Wire line needs to fall from VPUP below the threshold VTL. To get from active to idle, the voltage needs to rise from VIL- MAX past the threshold VTH. The time it takes for the

voltage to make this rise is seen in Figure 12 as “ε” and

its duration depends on the pullup resistor (RPUP) used

and the capacitance  of the 1-Wire network attached.

The voltage VILMAX is relevant for the DS1922E when determining a logical level, not triggering any events.

The initialization sequence required  to begin any com- munication with the DS1922E is shown in Figure 12. A reset pulse followed by a presence pulse indicates the DS1922E is ready to receive data, given the correct ROM and memory function command. If the bus master uses slew-rate control on the falling edge, it must pull down the line for tRSTL + tF to compensate for the edge. A tRSTL duration of 690µs or longer exits the Overdrive Mode, returning the device to standard speed. If the

DS1922E is in Overdrive Mode and tRSTL is no longer than 80µs, the device remains in Overdrive Mode.

After the bus master has released the line, it goes into receive mode (Rx). Now the 1-Wire bus is pulled  to VPUP through the pullup  resistor or, in the case of a DS2480B driver, through active circuitry.  When the threshold VTH is crossed,  the DS1922E waits for tPDH and then transmits a presence pulse by pulling the line low for tPDL. To detect  a presence  pulse, the master must test the logical state of the 1-Wire line at tMSP.

The t RSTH  window  must  be  at  least  the  sum  of tPDHMAX, tPDLMAX, and tRECMIN. Immediately  after tRSTH is expired,  the DS1922E is ready for data com- munication.  In a mixed  population  network, tRSTH should be extended  to minimum 480µs at standard speed  and 48µs at overdrive speed  to accommodate other 1-Wire devices.

Read/Write Time Slots Data communication with the DS1922E takes place in time slots that carry a single bit each. Write time slots transport data from bus master to slave. Read time slots transfer data from slave to master. The definitions of the write and read time slots are illustrated in Figure 13.

All communication  begins with the master pulling the data line low. As the voltage on the 1-Wire line falls below the threshold VTL, the DS1922E starts its internal timing generator that determines when the data line is sampled during a write time slot and how long data is valid during a read time slot.

MASTER Tx "RESET PULSE"	MASTER Rx "PRESENCE PULSE"

VPUP VIHMASTER VTH

ε	tMSP

VTL

VILMAX

0V

tF

tPDH

tRSTL 	tPDL

tRSTH

tREC

DS1922E

RESISTOR	MASTER	DS1922E

Figure 12. Initialization Procedure: Reset and Presence Pulse

Master-to-Slave For a write-one time slot, the voltage on the data line must have crossed the VTH threshold before the write-one low time tW1LMAX is expired. For a write-zero time slot, the voltage on the data line must stay below the VTH threshold until the write-zero low time tW0LMIN is expired. The voltage on the data line should not exceed VILMAX during  the entire tW0L or tW1L window. After the VTH threshold has been crossed,  the DS1922E needs a recovery time tREC before it is ready for the next time slot.

DS1922E

Slave-to-Master A read-data time slot begins like a write-one time slot. The voltage on the data line must remain below VTL until the read low time tRL is expired.  During the tRL window, when responding with a 0, the DS1922E starts pulling the data line low; its internal timing generator determines when this pulldown ends and the voltage starts rising again. When responding  with a 1, the DS1922E does not hold the data line low at all, and the voltage starts rising as soon as tRL is over.

The sum of tRL + δ (rise time) on one side and the inter- nal timing generator of the DS1922E on the other side define  the master sampling  window  (t MSRMIN  to tMSRMAX) in which the master must perform a read from the data line. For most reliable communication,  tRL should be as short as permissible  and the master should read close to but no later than tMSRMAX. After reading  from the data line, the master must wait until tSLOT is expired.  This guarantees sufficient recovery time tREC for the DS1922E to get ready for the next time slot. Note that tREC specified  herein applies only to a single DS1922E attached to a 1-Wire line. For multiple device  configurations,  tREC must be extended  to accommodate  the additional  1-Wire device  input

capacitance.  Alternatively, an interface that performs active pullup  during  the 1-Wire recovery time such as the DS2482-x00 or DS2480B 1-Wire line drivers can be used.

Improved Network Behavior

(Switchpoint Hysteresis)

In a 1-Wire environment line termination is possible only during transients controlled by the bus master (1-Wire driver). 1-Wire networks, therefore, are susceptible  to noise of various origins. Depending  on the physical size and topology  of the network, reflections from end

points and branch points can add up or cancel each other to some extent. Such reflections are visible as glitches  or ringing on the 1-Wire communication  line. Noise coupled  onto the 1-Wire line from external sources can also result in signal glitching. A glitch dur- ing the rising edge  of a time slot can cause a slave device to lose synchronization with the master and, as a consequence,  result in a Search ROM command coming to a dead end or cause a device-specific  func- tion command to abort. For better performance  in net- work applications,  the DS1922E uses a new 1-Wire front-end, which makes it less sensitive to noise and also reduces  the magnitude  of noise injected  by the slave device itself.

The DS1922E’s 1-Wire front-end differs from traditional slave devices in four characteristics:

1) The falling edge  of the presence pulse has a con- trolled slew rate. This provides a better match to the line impedance  than a digitally  switched transistor, converting  the high-frequency  ringing  known from traditional devices  into a smoother low-bandwidth transition. The slew-rate control is specified  by the parameter tFPD, which has different values for stan- dard and overdrive speed.

2) There is additional lowpass filtering in the circuit that detects the falling edge  at the beginning  of a time slot. This reduces  the sensitivity to high-frequency noise. This additional  filtering does not apply  at overdrive speed.

3) There is a hysteresis at the low-to-high switching threshold VTH. If a negative glitch crosses VTH but does not go below VTH - VHY, it is not recognized (Figure 14, Case A). The hysteresis is effective at any 1-Wire speed.

4) There is a time window specified by the rising edge hold-off time tREH during which glitches are ignored, even if they extend below VTH - VHY threshold (Figure 14, Case B, tGL <  tREH). Deep voltage droops or glitches that appear late after crossing the VTH threshold and extend beyond the tREH window cannot be filtered out and are taken as the begin- ning of a new time slot (Figure 14, Case C, tGL ≥ tREH).

Devices that have the parameters tFPD, VHY, and tREH specified  in their electrical  characteristics  use the improved 1-Wire front-end.

WRITE-ONE TIME SLOT

VPUP VIHMASTER VTH

tW1L

VTL

VILMAX

tF 	ε

tSLOT

RESISTOR	MASTER

WRITE-ZERO TIME SLOT

VPUP VIHMASTER VTH

tW0L

VTL

VILMAX

0V

tF

tSLOT

ε	tREC

RESISTOR	MASTER

READ-DATA TIME SLOT

VPUP VIHMASTER VTH

VTL

VILMAX

0V

tRL

tMSR

MASTER SAMPLING WINDOW

tF 	δ

tSLOT

tREC

DS1922E

RESISTOR	MASTER	DS1922E

Figure 13. Read/Write Timing Diagrams

| VPUP | tREH |  |  | tREH |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| VHY |  |  |  |  |  |  |
| CASE A |  | CASE B |  |  |  | CASE C |
|  | tGL |  |  | tGL |  |  |

VTH

0V

Figure 14. Noise Suppression Scheme

POLYNOMIAL = X16 + X15 + X2 + 1

1ST STAGE

2ND STAGE

3RD STAGE

4TH STAGE

5TH STAGE

6TH STAGE

7TH STAGE

8TH STAGE

X0 	X1 	X2 	X3 	X4

X5 	X6 	X7

9TH STAGE

10TH STAGE

11TH STAGE

12TH STAGE

13TH STAGE

14TH STAGE

15TH STAGE

16TH STAGE

X8 	X9 	X10

X11 	X12

X13 	X14

X15

X16

CRC OUTPUT

INPUT DATA

Figure 15. CRC-16 Hardware Description and Polynomial

CRC Generation

DS1922E

The DS1922E uses two types of CRCs. One CRC is an

8-bit type and is stored in the most significant byte of the 64-bit ROM. The bus master can compute  a CRC value from the first 56 bits of the 64-bit ROM and com- pare it to the value stored within the DS1922E to deter- mine if the ROM data has been received error-free. The equivalent polynomial function of this CRC is X8 + X5 + X4 + 1. This 8-bit CRC is received in the true (noninvert- ed) form, and it is computed at the factory and lasered into the ROM.

The other CRC is a 16-bit type, generated according  to the standardized CRC-16 polynomial function x16 + x15

+ x2  + 1. This CRC is used for error detection  when reading  register pages  or the data-log  memory using the Read Memory with CRC command and for fast veri- fication of a data transfer when writing to or reading from the scratchpad.  In contrast to the 8-bit CRC, the

16-bit CRC is always communicated  in the inverted form. A CRC generator inside the DS1922E (Figure 15) calculates a new 16-bit CRC as shown in the command flowchart of Figure 9. The bus master compares  the CRC value read from the device to the one it calculates from the data and decides whether to continue with an operation or to reread the portion of the data with the CRC error. With the initial pass through the Read Memory with CRC flowchart, the 16-bit CRC value is the result of shifting the command  byte into the cleared CRC generator, followed by the two address bytes and the data bytes. The password  is excluded  from the CRC calculation. Subsequent passes through the Read Memory with CRC flowchart generate a 16-bit CRC that is the result of clearing  the CRC generator and then shifting in the data bytes.

With the Write Scratchpad command, the CRC is gener- ated by first clearing the CRC generator and then shift-

ing in the command  code, the target addresses  TA1 and TA2, and all the data bytes. The DS1922E transmits this CRC only if the data bytes written to the scratchpad include scratchpad  ending offset 11111b. The data can start at any location within the scratchpad.

With the Read Scratchpad  command, the CRC is gen- erated by first clearing  the CRC generator and then

shifting in the command  code,  the target addresses TA1 and TA2, the E/S byte, and the scratchpad  data starting at the target address. The DS1922E transmits this CRC only if the reading continues through the end of the scratchpad,  regardless of the actual ending off- set. For more information on generating  CRC values, refer to Application Note 27.

DS1922E

Command-Specific 1-Wire Communication Protocol—Legend

| SYMBOL | DESCRIPTION |
| --- | --- |
| RST | 1-Wire reset pulse generated by master. |
| PD | 1-Wire presence pulse generated by slave. |
| Select | Command and data to satisf y the ROM function protocol. |
| WS | Command "Write Scratchpad." |
| RS | Command "Read Scratchpad." |
| CPS | Command "Copy Scratchpad with Password." |
| RMC | Command "Read Memory with Password and CRC." |
| CM | Command "Clear Memory with Password." |
| FC | Command "Forced Conversion." |
| SM | Command "Start Miss ion w ith Password." |
| STP | Command "Stop Miss ion w ith Password." |
| TA | Target Address TA1, TA2. |
| TA–E/S | Target Address TA1, TA2 with E/S byte. |
| <Data to EOS> | Transfer of as many data bytes as are needed to reach the scratchpad offset 1Fh. |
| <Data to EOP> | Transfer of as many data bytes as are needed to reach the end of a memory page. |
| <PW/Dummy> | Transfer of 8 bytes that either represent a valid password or acceptable dummy data. |
| <32 Bytes> | Transfer of 32 bytes. |
| <Data> | Transfer of an undetermined amount of data. |
| FFh | Transmis sion of one FFh byte. |
| CRC-16 | Transfer of an inverted CRC-16. |
| FF Loop | Indefinite loop where the master reads FF bytes. |
| AA Loop | Indefinite loop where the master reads AA bytes. |

Command-Specific 1-Wire Communication Protocol—Color Codes

Master-to-Slave 	Slave-to-Master

1-Wire Communication Examples

Write Scratchpad, Reaching the End of the Scratchpad (Cannot Fail)

| RST | PD | Select | WS | TA | <Data to EOS> | CRC-16 | FF Loop |
| --- | --- | --- | --- | --- | --- | --- | --- |

Read Scratchpad (Cannot Fail)

| RST | PD | Select | RS | TA-E/S | <Data to EOS> | CRC-16 | FF Loop |
| --- | --- | --- | --- | --- | --- | --- | --- |

Copy Scratchpad with Password (Success)

| RST | PD | Select | CPS | TA-E/S | <PW/Dummy> | AA Loop |
| --- | --- | --- | --- | --- | --- | --- |

Copy Scratchpad with Password (Fail TA-E/S or Password)

| RST | PD | Select | CPS | TA-E/S | <PW/Dummy> | FF Loop |
| --- | --- | --- | --- | --- | --- | --- |

Read Memory with Password and CRC (Success)

RST  PD

Select

RMC   TA

<PW/Dummy>     <Data to EOP>

CRC-16

<32 Bytes>

CRC-16

FF Loop

Read Memory with Password and CRC (Fail Password or Address)

Loop

DS1922E

| RST | PD | Select | RMC | TA | <PW/Dummy> | FF Loop |
| --- | --- | --- | --- | --- | --- | --- |

Clear Memory with Password

| RST | PD | Select | CM | <PW/Dummy> | FFh | FF Loop |
| --- | --- | --- | --- | --- | --- | --- |

To verify success, read the General Status register at address  0215h. If MEMCLR is 1, the command  was executed successfully.

1-Wire Communication Examples (continued)

Forced Conversion

| RST | PD | Select | FC | FFh | FF Loop |
| --- | --- | --- | --- | --- | --- |

To read the result and to verify success, read the addresses 020Ch to 020Fh (results) and the device samples counter at address 0223h to 0225h. If the count has incremented, the command was executed successfully.

Start Mission with Password

| RST | PD | Select | SM | <PW/Dummy> | FFh | FF Loop |
| --- | --- | --- | --- | --- | --- | --- |

To verify success, read the General Status register at address 0215h. If MIP is 1 and MEMCLR is 0, the command was executed successfully.

Stop Mission with Password

| RST | PD | Select | STP | <PW/Dummy> | FFh | FF Loop |
| --- | --- | --- | --- | --- | --- | --- |

To verify success, read the General Status register at address 0215h. If MIP is 0, the command was executed successfully.

Mission Example: Prepare and Start a New Mission

Assumption: The previous mission has been ended by using the Stop Mission command.  Passwords are not enabled.

Starting a mission requires three steps:

Step 1: Clear the data of the previous mission. Step 2: Write the setup data to register page 1. Step 3: Start the new mission.

Step 1: Clear the data of the previous mission.

With only a single device connected to the bus master, the communication of step 1 looks like this:

DS1922E

| MASTER MODE | DATA (LSB FIRST) | COMMENTS |
| --- | --- | --- |
| Tx | (Reset) | Reset pulse |
| Rx | (Presence) | Presence pulse |
| Tx | CCh | Issue “Skip ROM” command |
| Tx | 96h | Issue “Clear Memory” command |
| Tx | <8 FFh bytes> | Send dummy password |
| Tx | FFh | Send dummy byte |
| Tx | (Reset) | Reset pulse |
| Rx | (Presence) | Presence pulse |

Step 2: Write the setup data to register page 1.

During the setup, the device needs to learn the follow- ing information:

• Time and Date

• Sample Rate

• Alarm Thresholds

• Alarm Controls (Response to Conditional Search)

• General Mission Parameters (e.g., Channels to Log and Logging Format, Rollover, Start Mode)

• Mission Start Delay

The following data sets up the DS1922E for a mission that logs temperature using 8-bit format.

DS1922E

| ADDRESS | DATA | EXAMPLE VALUES | FUNCTION |
| --- | --- | --- | --- |
| 0200h | 00h | 15:30:00 hours | Time |
| 0201h | 30h |  |  |
| 0202h | 15h |  |  |
| 0203h | 01h | 1st of April in 2008 | Date |
| 0204h | 04h |  |  |
| 0205h | 08h |  |  |
| 0206h | 0Ah | Every 10 minutes (EHSS = 0) | Sample rate |
| 0207h | 00h |  |  |
| 0208h | 08h | 18°C low 135°C high | Temperature alarm thresholds |
| 0209h | F2h |  |  |
| 020Ah | 00h | (Don’t care) | (Not applicable w ith DS1922E) |
| 020Bh | FFh |  |  |
| 020Ch | FFh | (Don’t care) | Clock through read-only registers |
| 020Dh | FFh |  |  |
| 020Eh | FFh |  |  |
| 020Fh | FFh |  |  |
| 0210h | 02h | Enable high alarm | Temperature alarm control |
| 0211h | FCh | Disabled | (Not applicable w ith DS1922E) |
| 0212h | 01h | On (enabled), EHSS = 0 (low sample rate) | RTC oscillator control, sample rate select ion |
| 0213h | C1h | Normal start; no rollover; 8-bit temperature log | General mission control |
| 0214h | FFh | (Don’t care) | Clock through read-only registers |
| 0215h | FFh |  |  |
| 0216h | 5Ah | 90 minutes | Mis sion start delay |
| 0217h | 00h |  |  |
| 0218h | 00h |  |  |

With only a single device connected to the bus master, the communication of step 2 looks like this:

| MASTER MODE | DATA (LSB FIRST) | COMMENTS |  |
| --- | --- | --- | --- |
| Tx | (Reset) | Reset pulse |  |
| Rx | (Presence) | Presence pulse |  |
| Tx | CCh | Issue “Skip ROM” command |  |
| Tx | 0Fh | Issue “Write Scratchpad” command |  |
| Tx | 00h | TA1, beginning offset = 00h |  |
| Tx | 02h | TA2, address = 0200h |  |
| Tx | <25 Data Bytes> | Write 25 bytes of data to scratchpad |  |
| Tx | <7 FFh Bytes> | Write through the end of the scratchpad |  |
| Tx | (Reset) | Reset pulse |  |
| Rx | (Presence) | Presence pulse |  |
| Tx | CCh | Issue “Skip ROM” command |  |
| Tx | AAh | Issue “Read Scratchpad” command |  |
| Rx | 00h | Read TA1, beginning off set = 00h |  |
| Rx | 02h | Read TA2, address = 0200h |  |
| Rx | 1Fh | Read E/S, ending offset = 1Fh, flags = 0h |  |
| Rx | <32 Data Bytes> | Read scratchpad data and verify |  |
| Tx | (Reset) | Reset pulse |  |
| Rx | (Presence) | Presence pulse |  |
| Tx | CCh | Issue “Skip ROM” command |  |
| Tx | 99h | Issue “Copy Scratchpad” command |  |
| Tx | 00h | TA1 | (AUTHORIZATION CODE) |
| Tx | 02h | TA2 |  |
| Tx | 1Fh | E/S |  |
| Tx | <8 FFh Bytes> | Send dummy password |  |
| Tx | (Reset) | Reset pulse |  |
| Rx | (Presence) | Presence pulse |  |

Step 3: Start the new mission.

With only a single device connected to the bus master, the communication of step 3 looks like this:

If step 3 was successful,  the MIP bit in the General Status register is 1, the MEMCLR bit is 0, and the mis- sion start delay counts down.

DS1922E

| MASTER MODE | DATA (LSB FIRST) | COMMENTS |
| --- | --- | --- |
| Tx | (Reset) | Reset pulse |
| Rx | (Presence) | Presence pulse |
| Tx | CCh | Issue “Skip ROM” command |
| Tx | CCh | Issue “Start Mission” command |
| Tx | <8 FFh Bytes> | Send dummy password |
| Tx | FFh | Send dummy byte |
| Tx | (Reset) | Reset pulse |
| Rx | (Presence) | Presence pulse |

Software Correction Algorithm for Temperature

The correction  algorithm described  in the DS1922L/

Pin  Configuration

F5 SIZE

5.89mm

DS1922T data sheet does not apply to the DS1922E. If

attempted,  the corrected  result is generally less accu- rate than the raw temperature  data read from the device. Therefore, with the DS1922E the memory pages

18 and 19 are available as additional user memory.

0.51mm

A1

BRANDING

®   41

16.25mm

1-Wire®

Thermochron®

17.35mm

DS1922E

IO

GND

Package Information

For the latest package outline information and land patterns, go to  www.maxim-ic.com/packages.

| PACKAGE TYPE | PACKAGE CODE | DOCUMENT NO. |
| --- | --- | --- |
| F5 iButton | IB#6CB | 21-0266 |

DS1922E

Revision History

| REVISION NUMBER | REVISION DATE | DESCRIPTION | PAGES CHANGED |
| --- | --- | --- | --- |
| 0 | 7/08 | Init ial release. | — |
| 1 | 10/08 | Added the Software Correction Algorithm for Temperature section. | 43 |
| 2 | 6/09 | Changed storage temperature range in Absolute Maximum Rat ings and added a recommended storage temperature note for maximum battery lifetime. | 2 |

Maxim cannot assume responsibility  for use of any circuitry  other than circuitry  entirely embodied  in a Maxim product.  No circuit  patent licenses  are implied. Maxim reserves the right to change the circuitry and specifications  without notice at any time.

44  	Maxim Integrated Products, 120  San  Gabriel Drive, Sunnyvale, CA  94086 408-737-7600

©  2009 Maxim Integrated  Products 	Maxim is a registered  trademark  of Maxim Integrated  Products, Inc.