19-2235; Rev 1; 3/02

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

General Description

The MAX6675 performs cold-junction  compensation and digitizes  the signal from a type-K thermocouple. The data is output in a 12-bit resolution, SPI™-compati- ble, read-only format.

This converter resolves temperatures to 0.25°C, allows readings as high as +1024°C, and exhibits thermocou- ple accuracy  of 8LSBs for temperatures ranging  from

0°C to +700°C.

The MAX6675 is available in a small, 8-pin SO package.

MAX6675

Features

♦ Direct Digital Conversion of Type -K Thermocouple Output

♦ Cold-Junction Compensation

♦ Simple SPI-Compatible Serial Interface

♦ 12-Bit, 0.25°C Resolution

♦ Open Thermocouple Detection

Ordering Information

Applications

Industrial Appliances HVAC

PART                  TEMP RANGE          PIN-PACKAGE

MAX6675ISA               -20°C to +85°C           8 SO

Pin  Configuration

TOP VIEW

Automotive

| GND T- T+ | MAX6675 |  |
| --- | --- | --- |
| 1 |  | 8 |
| 2 |  | 7 |
| 3 |  | 6 |
| 4 |  | 5 |

VCC

N.C. SO CS

SCK

SO

SPI is a trademark of Motorola, Inc.

Typical Application Circuit

Vcc

0.1µF

GND

MAX6675

MICROCONTROLLER

68HC11A8

SO                        MISO

T+                                                        SCK

T-                                             CS

SCK

SSB

Maxim Integrated Products   1

For pricing, delivery, and ordering information, please contact Maxim/Dallas Direct! at

1-888-629-4642, or visit Maxim’s website at www.maxim-ic.com.

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

MAX6675

ABSOLUTE MAXIMUM RATINGS

Supply Voltage (VCC to GND)  ................................ -0.3V to +6V SO, SCK, CS, T-, T+ to GND  .......................-0.3V to VCC + 0.3V SO Current  ........................................................................ 50mA ESD Protection (Human Body Model)  ........................... ±2000V Continuous Power Dissipation  (TA = +70°C)

8-Pin SO (derate 5.88mW/°C above +70°C)  .............. 471mW

Operating  Temperature  Range  ..........................-20°C to +85°C

Storage Temperature  Range  ............................-65°C to +150°C Junction Temperature  .................................................... +150°C SO Package

Vapor Phase (60s) . .....................................................+215°C Infrared (15s)  ..............................................................+220°C Lead Temperature  (soldering,  10s)  ............................... +300°C

Stresses beyond those listed under “Absolute Maximum Ratings” may cause permanent damage to the device. These are stress ratings only, and functional operation of the device at these or any other conditions  beyond those indicated  in the operational sections of the specifications  is not implied. Exposure to absolute maximum rating conditions for extended periods may affect device reliability.

ELECTRICAL CHARACTERISTICS

(VCC = +3.0V to +5.5V, TA = -20°C to +85°C, unless otherwise noted. Typical values specified at +25°C.) (Note 1)

| PARAMETER | SYMBOL | CONDITIONS |  |  | MIN TYP MAX | UNITS |
| --- | --- | --- | --- | --- | --- | --- |
| Temperature Error |  | TTHERMOCOUPLE = +700°C, TA = +25°C (Note 2) |  | VCC = +3.3V | -5 +5 | LSB |
|  |  |  |  | VCC = +5V | -6 +6 |  |
|  |  | TTHERMOCOUPLE = 0°C to +700°C, TA = +25°C (Note 2) |  | VCC = +3.3V | -8 +8 |  |
|  |  |  |  | VCC = +5V | -9 +9 |  |
|  |  | TTHERMOCOUPLE = +700°C to +1000°C, TA = +25°C (Note 2) |  | VCC = +3.3V | -17 +17 |  |
|  |  |  |  | VCC = +5V | -19 +19 |  |
| Thermocouple Conversion Constant |  |  |  |  | 10.25 | µV/LSB |
| Cold-Junction Compensation Error |  | TA = -20°C to +85°C (Note 2) | VCC = +3.3V |  | -3.0 +3.0 | °C |
|  |  |  | VCC = +5V |  | -3.0 +3.0 |  |
| Resolution |  |  |  |  | 0.25 | °C |
| Thermocouple Input Impedance |  |  |  |  | 60 | kΩ |
| Supply Voltage | VCC |  |  |  | 3.0 5.5 | V |
| Supply Current | ICC |  |  |  | 0.7 1.5 | mA |
| Power-On Reset Threshold |  | VCC rising |  |  | 1 2 2.5 | V |
| Power-On Reset Hysteresis |  |  |  |  | 50 | mV |
| Conversion Time |  | (Note 2) |  |  | 0.17 0.22 | s |
| SERIAL INTERFACE |  |  |  |  |  |  |
| Input Low Voltage | VIL |  |  |  | 0.3 x VCC | V |
| Input High Voltage | VIH |  |  |  | 0.7 x VCC | V |
| Input Leakage Current | ILEAK | VIN = GND or VCC |  |  | ±5 | µA |
| Input Capacitance | CIN |  |  |  | 5 | pF |

2

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

MAX6675

ELECTRICAL CHARACTERISTICS (continued)

(VCC = +3.0V to +5.5V, TA = -20°C to +85°C, unless otherwise noted. Typical values specified at +25°C.) (Note 1)

| PARAMETER | SYMBOL | CONDITIONS | MIN TYP MAX | UNITS |
| --- | --- | --- | --- | --- |
| Output High Voltage | VOH | ISOURCE = 1.6mA | VCC - 0.4 | V |
| Output Low Voltage | VOL | ISINK = 1.6mA | 0.4 | V |
| TIMING |  |  |  |  |
| Serial Clock Frequency | fSCL |  | 4.3 | MHz |
| SCK Pulse High Width | tCH |  | 100 | ns |
| SCK Pulse Low Width | tCL |  | 100 | ns |
| CSB Fall to SCK Rise | tCSS | CL = 10pF | 100 | ns |
| CSB Fall to Output Enable | tDV | CL = 10pF | 100 | ns |
| CSB Rise to Output Disable | tTR | CL = 10pF | 100 | ns |
| SCK Fall to Output Data Valid | tDO | CL = 10pF | 100 | ns |

Note 1: All specifications are 100% tested at TA = +25°C. Specification limits over temperature  (TA = TMIN to TMAX) are guaranteed by design and characterization, not production tested.

Note 2: Guaranteed  by design. Not production tested.

Typical Operating Characteristics

(VCC = +3.3V, TA = +25°C, unless otherwise noted.)

OUTPUT CODE ERROR

MAX6675 toc01

vs. AMBIENT TEMPERATURE

10

OUTPUT  CODE ERROR (LSB)

8

6

4

2

OUTPUT CODE ERROR

MAX6675 toc02

vs. VOLTAGE DIFFERENTIAL

10

5

OUTPUT  CODE ERROR (LSB)

0

0

| 0 15 30 45 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

60       75       90

-5

| -10 0 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

10       20       30

40       50

TEMPERATURE (°C)

VOLTAGE DIFFERENTIAL (mV)

3

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

MAX6675

Pin  Description

| PIN | NAME | FUNCTION |
| --- | --- | --- |
| 1 | GND | Ground |
| 2 | T- | Alumel Lead of Type-K Thermocouple. Should be connected to ground externally. |
| 3 | T+ | Chromel Lead of Type-K Thermocouple |
| 4 | VCC | Positive Supply. Bypass with a 0.1µF capacitor to GND. |
| 5 | SCK | Serial Clock Input |
| 6 | CS | Chip Select. Set CS low to enable the serial interface. |
| 7 | SO | Serial Data Output |
| 8 | N.C. | No Connection |

Detailed Description

The MAX6675 is a sophisticated  thermocouple-to-digi- tal converter with a built-in 12-bit analog-to-digital  con- verter (ADC). The MAX6675 also contains cold-junction compensation  sensing and correction,  a digital  con- troller, an SPI-compatible interface, and associated control logic.

The MAX6675 is designed to work in conjunction with an external microcontroller (µC) or other intelligence in ther- mostatic, process-control, or monitoring applications.

Temperature Conversion The MAX6675 includes signal-conditioning  hardware to convert the thermocouple’s signal into a voltage compat- ible with the input channels of the ADC. The T+ and T- inputs connect to internal circuitry that reduces the intro- duction of noise errors from the thermocouple wires.

Before converting  the thermoelectric  voltages into equivalent temperature values, it is necessary to com- pensate for the difference  between the thermocouple cold-junction side (MAX6675 ambient temperature) and a 0°C virtual reference. For a type-K thermocouple, the voltage changes  by 41µV/°C, which approximates the thermocouple  characteristic  with the following linear equation:

VOUT = (41µV / °C) ✕ (TR - TAMB)

Where:

VOUT is the thermocouple output voltage (µV).

TR is the temperature of the remote thermocouple junc- tion (°C).

TAMB is the ambient temperature (°C).

Cold-Junction Compensation The function of the thermocouple is to sense a differ- ence in temperature between two ends of the thermo- couple wires. The thermocouple‘s hot junction can be read from 0°C to +1023.75°C. The cold end (ambient temperature  of the board  on which the MAX6675 is mounted) can only range from -20°C to +85°C. While the temperature  at the cold  end  fluctuates,  the MAX6675 continues to accurately  sense the tempera- ture difference at the opposite end.

The MAX6675 senses and corrects for the changes in the ambient temperature with cold-junction  compensa- tion. The device  converts the ambient temperature reading  into a voltage using a temperature-sensing diode.  To make the actual thermocouple temperature measurement, the MAX6675 measures the voltage from the thermocouple’s output and from the sensing diode. The device’s internal circuitry  passes the diode’s volt- age (sensing ambient temperature) and thermocouple voltage (sensing remote temperature minus ambient temperature) to the conversion function stored in the ADC to calculate the thermocouple’s hot-junction tem- perature.

Optimal performance  from the MAX6675 is achieved when the thermocouple cold junction and the MAX6675 are at the same temperature. Avoid placing heat-gener- ating  devices  or components  near the MAX6675 because this may produce cold-junction-related errors.

Digitization The ADC adds  the cold-junction  diode  measurement with the amplified thermocouple voltage and reads out the 12-bit result onto the SO pin. A sequence  of all zeros means the thermocouple  reading  is 0°C. A sequence of all ones means the thermocouple reading is +1023.75°C.

4

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

Applications Information

Serial Interface The Typical Application  Circuit shows the MAX6675 interfaced  with a microcontroller.  In this example, the MAX6675 processes  the reading  from the thermocou- ple and transmits the data through a serial interface. Force CS low and apply a clock signal at SCK to read the results at SO. Forcing CS low immediately stops any conversion  process.  Initiate a new conversion process by forcing CS high.

Force CS low to output the first bit on the SO pin. A complete serial interface read requires 16 clock cycles. Read the 16 output bits on the falling edge of the clock. The first bit, D15, is a dummy sign bit and is always zero. Bits D14–D3 contain the converted temperature in the order of MSB to LSB. Bit D2 is normally low and goes high when the thermocouple input is open. D1 is low to provide a device ID for the MAX6675 and bit D0 is three-state.

Figure 1a is the serial interface protocol and Figure 1b shows the serial interface timing. Figure 2 is the SO out- put.

Open Thermocouple Bit D2 is normally low and goes high if the thermocou- ple input is open. In order to allow the operation of the open thermocouple  detector,  T- must be grounded. Make the ground connection as close to the GND pin as possible.

Noise Considerations The accuracy of the MAX6675 is susceptible to power- supply  coupled  noise. The effects of power-supply noise can be minimized by placing  a 0.1µF ceramic bypass capacitor close to the supply pin of the device.

Thermal Considerations Self-heating degrades  the temperature measurement accuracy  of the MAX6675 in some applications.  The magnitude  of the temperature errors depends  on the thermal conductivity  of the MAX6675 package,  the

MAX6675

mounting technique,  and the effects of airflow. Use a large ground  plane to improve the temperature mea- surement accuracy of the MAX6675.

The accuracy  of a thermocouple  system can also be improved by following these precautions:

•  Use the largest wire possible  that does not shunt heat away from the measurement area.

•  If small wire is required, use it only in the region of the measurement and use extension wire for the region with no temperature gradient.

•  Avoid mechanical stress and vibration, which could strain the wires.

•  When using long thermocouple wires, use a twisted- pair extension wire.

•  Avoid steep temperature gradients.

•  Try to use the thermocouple wire well within its tem- perature rating.

•  Use the proper sheathing material in hostile environ- ments to protect the thermocouple wire.

•  Use extension wire only at low temperatures  and only in regions of small gradients.

•  Keep an event log and a continuous record of ther- mocouple resistance.

Reducing Effects of Pick-Up Noise The input  amplifier  (A1) is a low-noise amplifier designed to enable high-precision input sensing. Keep the thermocouple  and connecting  wires away from electrical noise sources.

Chip Information

TRANSISTOR COUNT: 6720

PROCESS: BiCMOS

5

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

CS

MAX6675

SCK

SO                                                                                                                                                                                                                                                    D0

D15        D14        D13        D12        D11        D10         D9          D8          D7          D6          D5          D4          D3          D2          D1

Figure 1a. Serial Interface Protocol

tCSS

CS

tCH                                                                  tCL

SCK

tDV

tDO

tTR

SO

D15

D3         D2       D1        D0

Figure 1b. Serial Interface Timing

| BIT | DUMMY SIGN BIT | 12-BIT TEMPERATURE READING |  |  |  |  |  |  |  |  |  |  |  | THERMOCOUPLE INPUT | DEVICE ID | STATE |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bit | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
|  | 0 | MSB |  |  |  |  |  |  |  |  |  |  | LSB |  | 0 | Three- state |

Figure 2. SO Output

6

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

MAX6675

Block Diagram

VCC

0.1µF

4

DIGITAL CONTROLLER

COLD-JUNCTION

COMPENSATION                             S5

DIODE

5

SCK

300kΩ

T+         S3

3

T-

30kΩ

S2

30kΩ

S1

1MΩ

A1                                                                                             A2

S4

ADC               7

SO

2                                                                                                                                                                                                                               6

20pF                                                                                                                                       CS

300kΩ

MAX6675

REFERENCE VOLTAGE

1

GND

7

Cold-Junction-Compensated K-Thermocouple- to-Digital Converter (0°C  to +1024°C)

MAX6675

SOICN.EPS

Package Information

Maxim cannot assume responsibility  for use of any circuitry  other than circuitry  entirely embodied  in a Maxim product.  No circuit  patent licenses  are implied. Maxim reserves the right to change the circuitry and specifications  without notice at any time.

8                                       Maxim Integrated Products, 120  San  Gabriel Drive, Sunnyvale, CA  94086 408-737-7600

©  2002 Maxim Integrated  Products               Printed USA                                  is a registered  trademark  of Maxim Integrated  Products.